Laboratory of hardware for non-destructive testing (mirea)

Software Diada 3.16

(Description and manual)

18.12.2012

Diada - is software to acquire, visualization and digital processing of 16 bits per pixel half-tone images and its sequences. All parts of the instruction manual and the software are subject to the copyright. All rights are protected. Any duplication completely or partially is strictly forbidden. No part of the documentation and software may be copied, reproduced and translated or stored on another electronic medium or may be brought to mechanically readable form. To do this, in any case get in contact with TSNK-lab of MIREA.

Contents

First Steps	5

Install Program	5

Uninstall Program	5

Run Program	5

Get Images	5

Read Files	5

Read binary files	6

Collect Files to Sequence	7

Load Floating Point Data from Separate Files	7

Work with Video-Camera	7

Setup Video-Camera	8

Frame Acquisition by Video-Camera	9

Frame Sequence Acquisition by Video-Camera	10

Work with digital devices	10

tsnkCAM - camera	10

SensiCam-camera	12

PaxScan4030	12

PaxScan2520	13

Photo-camera Canon	14

Photo-camera Olympus	15

Hamamatsu - detector	16

Scanner Duerr NDT	17

DRTECH - panel (Korea)	17

VaTech panel (Korea)	18

Select Parameters for Image Acquisition	19

Work with TWAIN-compatible Device	19

Select TWAIN-source	19

Image Acquisition by TWAIN-protocol	20

Setup Duerr-NDT scanner	20

Scanning Systems to Acquire Images	20

X-ray Scanning System <Orel>	20

Prepare Scanning	20

Creation and Marking of Panoramic View	22

Frames of Region of Panoramic View	23

Tomosyntez of Region	24

X-ray Scanning System <Compton-Reverse>	24

X-ray Control Systems	25

Image Visualization Methods	28

Half-tone Image	29

Surface View	29

Tomogram View	29

Insight View	30

3D-View	31

Change Image Scale	31

Use Two Displays to Image Scaling	31

Scroll through a Big Images	32

Use Pseudo-colors Palettes	32

Use <Look Up and Histogram> tool	33

Use <Color Table> tool	33

Image Sewing from Frames	34

Measurements	35

Calibration of the Pixel Values, Length, Time and Angles	35

Length Calibration by Standard	35

Pixel Value Measurement	36

Pixel Values along Lines (Scans)	36

Image Distance Measurement	37

Minimal, Mean, Maximal and Dispersion Pixel Value Measurements For the Rectangular Region	37

Perimeter, Square, Minimal, Mean and Maximal Pixel Value Measurements	37

Reference Point Definition	37

Set Marks on Image	37

Measurement angles and conjective distances	38

Statistics for Ring or Round Regions	38

Brightness Table	39

Sharpness Estimation	39

Estimation of Scattered Radiation	39

Detector Persistence	39

Statistics of Detectors	40

Digital Image Processing	40

Region of Interest Selection	40

Image Editing	41

Image Resize	41

Change the Number and Order of Frames	41

Image Crop	41

Image Gluing	41

Erase Image Region	42

Pixel Values Scaling	42

Image lines edit	42

Image Transformations	43

Image Rotation	43

Mirror Transformation	43

Orthogonal Transformation	44

Affin Transformation	44

Non-Linear Transformation	44

Transformation <Fish Eye>	44

Trapezoid transformation	44

Logarithm of Fourier Spectrum	45

Background Correction	45

Subtract Background Image	45

Subtract Calculated (Linear or Parabolic) Background	45

Black & White Correction	45

Image Equalization	45

Correction of the Linear Detector Images	46

Correction of blur	46

PaxScan Corrections	47

Optic Distortion Corrections	47

Digital Filers	48

Set Filter Sizes and Conditions	48

Non-rectangular Filter Definition	48

Mean	49

Rank Statistics Filters	49

Convolution	49

Image Sharpening	49

Background Equalization	50

Non-linear Filters	50

User Defined Filter	51

Frequency Filters	52

Image Sequence Processing	52

Frame Sequence Too Big for Memory	52

Image Arithmetic Operations	52

Subtract Frame from Sequence	53

Average Frame Sequence	53

Time Averaging	53

Averaging with Horizontal Shift	53

Space Alignment	54

Image Analysis	54

Image Characteristics	54

Image Comparison	54

Image Binarization	54

Contours Extraction	54

Recognition of the point objects	54

BGA testing	55

Material's Recognition by dual energy method	56

Dual Energy Method Calibration	56

Acquire Two Images of Dual Energies	57

Metal and organics thicknesses estimation	57

Create a dual energy table	58

Tomosyntez	58

Use 2 Image Marks for Tomosyntez in Non-stationary Conditions	58

Use 4 Image Marks for Tomosyntez in Non-stationary Conditions	59

Correction of the Reconstructed Layers for the Cone Beam	59

CT - reconstruction	59

2D CT for the Classical Projection Data	59

Pseudo 3D-CT	60

Prepare Documentation	61

Drawing on Images	61

Save Image	62

Copy Images to Clipboard	62

Print Image	62

Create Demo Movies	62

Customize Program to User	63

Customize User Interface	63

Customize Commands	63

Customize Menu	64

Customize Toolbars	64

Add User Tools	65

Customize Keyboard Keys	65

Customize Double Mouse Click	66

Customize Styles of Toolbars and Menus	66

Create and Use Macro-Commands	66

Fool-Proof Protection	67

File Formats	67

First Steps

## Install Program

1. Switch off computer.

2. Insert in USB slot of computer the Hardlock key.

3. Switch on computer.

4. Run program hldrv32.exe from the hardlock driver disk. The newest version of the driver one can take from www.hardlock.com.

5. Run Setup.exe

6. Select English or Russian language for the installation and program-interface.

7. Insert serial number:

DEMO - for demonstration version,

12345 - for version with password (E5X4P3E2R1T) protection of the program settings. Where 12345 are symbols of the Hardlock, which you can see on the key or in file README.TXT.

8. After installation is finished run program by:

double click left mouse button on the desktop icon Diada,

or select menu Start -> Program -> Diada

## Uninstall Program

To uninstallation of the program - program select:

1. In the Task bar - Start - Settings - Control panel.

2. In the item "Install and uninstall software" select "Diada" and "Delete".

## Run Program

To run program you can:

double click the left mouse button on the desktop icon ,

or in the Task bar press Start - Programs - Diada – Diada

# Get Images

## Read Files

The command Open from File menu or click the button displays menu for loading of the existing image (sequence) from a file.

Click the button Preview for displaying of the image and information about its width, height and number bit on pixel.

The program supports the following pictorial formats of files:

AIM format (*.AIM)

AVI-format (*.AVI)

B16-format (*.B16)

BIN-format (*.BIN)

Bitmap Windows (*.BMP)

DTV-format (*.DTV)

Flexible image transport System format (*.FIT)

AutoDesk-format (*.FLC)

AutoDesk-format (*.FLI)

CompuServe GIF (*.GIF)

JPEG-format (*.JPG)

PC Paintbrush (*.PCX)

Portable Network Graphics (*.PNG)

Raw Binary Array (*.RAW)

Frame sequence (*.SET)

TIFF-format (*.TIF)

Text Table (*.TXT)

XRD-format (*.XRD)

XYZ-format (*.XYZ).

Notice that the program automatically converts the color images to the halftone ones, keeping the number of bits per pixel, if this number is not above 16.

|  | If the selected file with a unknown by expansion, then the additional dialog will appear, in which you can specify a type of a loaded file. |
| --- | --- |

After loading a file in the same folder it is searching a WMF-file with the same name as file have. If this file exists - it is loaded and the corresponding picture is shown on image.

## Read binary files

|  | For RAW-files there is a dialogue for input of the additional information about the size of the images, type of pixel values and quantities of bytes which are not containing to the useful information. Check the button Find minimal and maximal values if you want expand image pixels to range 0-65535. Check the button MAC-data format (byte inversion) to read data with inverse order of bytes which is usual for Macintosh-computers. Calculated file size is defined by all dialog's values and can be compared with true file size shown in dialog's title. |
| --- | --- |
|  | If one press button Subregion... there would be a dialog to define subregion in RAW-file. After press OK one return back to the original dialog with the correctly defined skip-values. |

|  | If one press the button Compress... there would be a dialog to set sizes of a frame sequence compressed from RAW-file data. All skipped bytes are not used in compression. The compression is produced by averaging old pixel values. This option is helpful if there is no enough memory to load a big RAW-file. |
| --- | --- |

## Collect Files to Sequence

The command sequence File, Operation, Collect Frames... or click button shows a dialog to create a frame sequence directly from separated files of frames.

Select frames one by one by left mouse button click. To delete frames you have to mark those frames and press button Remove frames.

After pressing button Create movie... in the current folder is created SET-file contained all selected frames, and a new created frame sequence is loading to a new window.

It is impossible to create such file on media, for example CD, which has read-only attribute.

## Load Floating Point Data from Separate Files

The command sequence File, Operation, Collect Data... or click button shows a dialog to create a frame sequence directly from separated files with the floating point data.

1. Press button Get from.. to select one of the data files. Program extracts the base filename from it automatically. Please, don't use "." symbol in file names. Use it only before filename extension.

2. Insert the width and height of the data files.

3. Define a start number of data file, increment (can be negative) and a total number of data files to load.

4. Press OK.

## Work with Video-Camera

Command Video-camera from File menu of click the button shows or hides a Video-window with the "live" image from a video camera. If on your computer more than one device for capture of a video is registered, the dialog for a select of the device will appear.

### Setup Video-Camera

To capture of the separate frame or sequence of the frames and setup of capture parameters - open the system menu of Video-window.

Live video - switch on or switch off a mode "Overlay" - the accelerated displaying of the image by its immediate transfer in a memory of the video-adapter, passing the buffer on a hard disk of the computer.

|  | Video source... - displays dialog to select a source and standard of video signal, and also Controls of brightness, contrast, saturation and hue. |
| --- | --- |

|  | Video-format... - displays dialog to select format and size of a video-frame. |
| --- | --- |

Select Region of Interest... - displays menu for a select of a working area of a video-window, which used for capture only of a subarea of all frame. You can select a subarea also with the help of the mouse with the pressed left button. This subarea is marked with red color.

### Frame Acquisition by Video-Camera

|  | Frame acquisition... - displays dialog to select the parameters of capture of the frame |
| --- | --- |

### Frame Sequence Acquisition by Video-Camera

|  | Create AVI-file... - displays dialog to select the parameters of capture of a sequence of the frames and their record into a AVI-file. |
| --- | --- |

## Work with digital devices

Before a work with the digital device it is necessary to select a type of the device by commands of the menu: File, Digital-device, Select device.... The dialog to select the device will appear.

1. Use this menu to select one from the supported digital cameras: tsnkCAM, SensiCam(PCO), PaxScan4030, PaxScan2520, PaxScan2520V, Hahamatsu, DRTECH, VaTech, photo-camera Canon, Olympus or scanner Duerr NDT.

2. Check the button Prompt to Exposure Time to get prompt of the exposure times and the number of digital accumulation cycles after frame acquisition. See also Select Parameters for Image Acquisition.

3. Check the button Auto switch X-ray radiation ON/OFF to operate by X-ray unit during image acquisition.

### tsnkCAM - camera

1. To setup tsnkCAM execute commands File, Digital-device, Setup... or click the button .

|  | 2. The following dialog to setup of camera parameters will appear. Those parameters can be selected only once before a several image acquisitions. 3. Select the number of Digital accumulation frames. 4. Define the Exposure Time. 5. Set Gain to maximize the brightness range of the acquired images. (This parameter is specific for each camera in multi-camera systems.) 6. Check the button Live video for continuous acquire images to one window. This mode is helpful during setup camera. 7. To accelerate frame rate (specially in live video mode) one can set binning equals to 2 or 4. 8. To set parameters quickly one can use buttons Save parameters... and Load parameters... |
| --- | --- |

|  | 9. Define the Delay before acquire. 10. Set Black level to shift the brightness range of the acquired images. (This parameter is specific for each camera in multi-camera systems.) 11. Set the camera Aperture. (Camera not supports this function.) 12. If there is a Automatic exposure adaptation mode, then the Exposure Time is defined automatically by camera. 13. Adaptation parameter is used by camera in the automatic adaptation mode to extrapolate the exposure time by the histogram of the testing image. 14. Parameter Percent of pixels to show is used by computer under the checked Normalization by histogram button to expand the dynamic range of the acquired image. |
| --- | --- |
|  | 15. Check button Cut region to crop a region from acquired image. 16. Press button Get from window to get from the active window a sizes of cutted region. 17. Select one of the image-processing operations from a list Post image processing. For Post-Scanner there must be selected the corresponding operation. 18. Check button Calibration by black image to acquire or/and load black calibration image. 19. Check button Calibration by white image to acquire or/and load white calibration image. |

20. Check button Optic distortion correction to acquire or/and load optic correction file. This button must be checked for the multi-camera systems and selected the corresponding optic correction file.

21. By Binning buttons one can compress the acquired image.

|  | 22. Set the number of cameras. 23. Select the active camera. The following image acquisition would be done by this camera only. The automatic image acquisition from several cameras is possible only during Parallel scanning for the System to Scan Lengthy Objects with Frame Sewing. For multi-tsnkCAM systems the button Matrix must be checked. In this case it is not important which camera is selected as active. In this case Lider camera is important. The lider camera receives acquisition command from computer and resends it to all other cameras. |
| --- | --- |

24. Set the network IP-address of camera. For multi-camera systems it is IP-address of first one. Press button Apply to set communication with camera and receive from it a camera's version and temperatures.

25. Press OK to send parameters from all pages of this dialog to the corresponding cameras.

To capture a frame by the tsnkCAM execute commands: File, Digital-device, Acquire or click the button .

### SensiCam-camera

|  | To setup SensiCam-camera execute commands File, Digital-device, Setup... or click the button . The following dialogs to setup of camera parameters will appear. |
| --- | --- |

For capture of the frame or the sequences of the frames by the SensiCam-camera execute commands: File, Digital device, Acquire or click the button .

### PaxScan4030

1. To setup PaxScan4030 execute commands File, Digital device, Setup... or click the button .

2. The following dialog to setup of PaxScan4030 parameters will appear. Those parameters can be selected only once before a several image acquisitions.

3. At first define correct orientation of the acquired image by Rotate radio-buttons and two Mirror-buttons.

4. Select ROI on full-size acquired image and press button Get ROI, to insert ROI sizes to dialog.

5. Press buttons Dark, Middle and Light, to load the calibration images from files, or to acquire new one.

6. Select which corrections you would like to apply.

Linear filter - subtract the background accumulated in first 64 columns of CCD-matrix.

PaxScan4030.err - correct the defects of CCD-matrix which are saved in a text file PaxScan4030.err by bilinear interpolation.

Dark subtract - subtract black calibration image.

Gain correction - makes a piecewise linear interpolation of brightness.

7. Select the number of digital averaging cycles.

8. Check button Flash-mode to set flash acquisition mode.

9. Check button Clean image to twice acquire each image. The first acquisition is used to clean CCD-matrix.

10. Parameters Bits per pixel (number of bits per pixel) and Dll name (name of PaxScan DLL-driver) must be selected in corresponding to your model of PaxScan. (See specific PaxScan-manual.)

11. Press OK.

To capture a frame by the PaxScan4030 Panel execute commands: File, Digital device, Acquire or click the button .

It is pity, that exposure time for PaxScan 4030 can be changed by hardware button only and it can't be more then 3 seconds.

At the same time the software FlashScan35 can set exposure times up to several minutes using undocumented functions of frame-grabber. Because there was realized a procedure of frame acquisition via FlashScan35 as puppet of this program. See below in Work with FlashScan35.

### PaxScan2520

1. To setup PaxScan4030 execute commands File, Digital device, Setup... or click the button .

2. The following dialog to setup of PaxScan2520 parameters will appear. Those parameters can be selected only once before a several image acquisitions.

3. At first define correct orientation of the acquired image by Rotate radio-buttons and two Mirror-buttons.

4. Select Model and Frame rate. The exposure time is back proportional to the frame rate.

5. Select ROI on full-size acquired image by mouse and press button Get ROI, to insert ROI sizes to dialog.

6. Use buttons Dark subtrack and Gain correction, to automatic creation of calibration images and save they in memory of command processor.

7. Select which corrections you would like to apply.

Linear filter - subtract the background accumulated in first columns of CCD-matrix.

Defect pixels - correct the defects of CCD-matrix which are saved in a text file PaxScan2520.err by bilinear interpolation.

Dark subtract - subtract black calibration image.

Gain correction - makes a piecewise linear interpolation of brightness.

8. Select the number of digital averaging cycles.

9. Check button Flash-mode to set flash acquisition mode.

10. Check button Clean image to twice acquire each image. The first acquisition is used to clean CCD-matrix.

11. Press OK.

To capture a frame by the PaxScan2520 Panel execute commands: File, Digital device, Acquire or click the button .

PaxScan2520V

#### PaxScan2520V

1. To setup PaxScan2520V execute commands File, Digital device, Setup... or click the button .

2. The following dialog to setup of PaxScan2520V parameters will appear. Those parameters can be selected only once before a several image acquisitions.

3. At first define correct orientation of the acquired image by Rotate radio-buttons and two Mirror-buttons.

4. Select Model, Frame frequency and number of the digital averaging cycles. The exposure time is back proportional to the frame frequency.

5. Select ROI on full-size acquired image by mouse and press button Get ROI, to insert ROI sizes to dialog.

6. Use buttons Offset correction and Gain correction, for automatic creation of calibration images and save they in the configuration folder.

7. Select which corrections you would like to apply.

Defect correction - correct the defects of CCD-matrix.

Black correction - subtract black calibration image.

Gain correction - makes a piecewise linear interpolation of brightness.

8. To test the image quality press button Snapshot.

9. To switch X-ray ON before acquire and switch OFF after acquire check the button Auto switch X-ray radiation ON/OFF in Select digital device dialog.

10. To get a live video window press button Live Video. To stop this mode use button Stop.

To capture a frame by the PaxScan2520V panel execute commands: File, Digital device, Acquire or click button .

### Photo-camera Canon

There are used a Canon firmware CD-SDK version 7.1. It must be installed during the program installation. The driver of Canon-camera itself must be installed separatelly. There were tested the following Canon models: PowerShot G2, PowerShot S70, PowerShot A620, A640.

Canon support button must be checked in Options dialog.

|  | 1. To setup Canon-camera execute commands File, Digital device, Setup... or click button . (If the Canon camera is not connected or its model is not a computer controllable one, nothing happens.) 2. There would be a dialog with live image from Canon camera. 3. Select zoom factor of picture by slider. The digital zoom is applied but not shown on viewfinder picture. 4. Press button AE/AF to get a sharp image. 5. To acquire a photo press button Release. |
| --- | --- |

|  | 6. The Camera type list select one of the connected to computer Canon cameras as active one. 7. Shooting mode - list of camera's working modes. 8. Size - list of frame sizes. 9. Quality - list of JPEG-format compressions. 10. Exposure is exposure time in seconds. 11. Diaphragm (Av) - list of Av - values. Big values means smaller open objective. 12. Flash - list of working modes of the internal flash. 13. By increasing ISO speed one can increase an image brightness even when a weak illumination conditions. 14. Check button Disconnect camera to automatic switch off camera after close this dialog and after one photo acquisition. |
| --- | --- |

To capture a frame by the Canon - camera without show the Setup-dialog execute commands: File, Digital device, Acquire or click the button .

### Photo-camera Olympus

There are used a Olympus firmware Mocha-SDK version 5.15. It must be installed during the program installation. The driver of Olympus-camera itself must be installed separately. There was tested the following Olympus model: E-420.

Olympus support button must be checked in Options dialog.

1. To setup Olympus-camera execute commands File, Digital device, Setup... or click the button .

2. There would be dialog parameters of Olympus camera. The parameters can be selected once before a sequence of acquisitions.

3. The Camera list select one of the connected to computer Olympus cameras as active one.

4. The Refresh button recreates a camera list.

5. To acquire a photo press button Capture.

6. Fix the AutoSave button to select a folder to save acquired frames and auto naming its.

7. The Clean Camera button removes all frames from camera memory.

8. Shooting mode - list of camera's working modes.

9. Resolution - list of frame sizes.

10. Focus Mode - defines mode of focusing.

11. Shutter speed is exposure time in seconds.

12. Aperture - list of aperture values. Big values means smaller open objective.

13. Flash - list of working modes of the internal flash.

14. By increasing ISO speed one can increase an image brightness even when a weak illumination conditions.

After camera OFF the camera settings are not saved and get the default ones.

|  | This page of dialog shows the parameters of Olympus camera: it's possible, default and current values. |
| --- | --- |
|  | This page of dialog shows the protocol of executed commands for Olympus camera. |

To capture a frame by the Olympus - camera without show the Setup-dialog execute commands: File, Digital device, Acquire or click the button .

### Hamamatsu - detector

To work with Hamamatsu flat-panels and linear detectors there is used a frame-grabber driver. It must be installed previously and separately from the program installation.

1. To setup Hamamatsu-detector execute commands File, Digital device, Setup... or click the button .

|  | 2. The following dialog to setup parameters will appear. Those parameters can be selected only once before a several image acquisitions. 3. Select the configuration file of your Hamamatsu device. 4. Input cutted region coordinates or select ROI by mouse and press button Get ROI. 5. Set the number of averaged frames. 6. Press button OK. |
| --- | --- |

### Scanner Duerr NDT

Typically to acquire an image from scanner one use TWAIN-drive. But for the Duerr NDT scanner there is an alternative method of INI-files. To use it:

1. Select Duerr NDT as current digital acquisition device.

2. Execute commands File, CCD-device, Setup... or click the button .

|  | 3. In the popup dialog one can set Resolution of scanning and a name of PNG-file to save result data. If this name is empty, a new image would be loaded but not be saved. 4. Press button Scan. |
| --- | --- |
|  | 5. This popup firmware dialog is shown while the image acquisition process. |

### DRTECH - panel (Korea)

DRTECH-panel can be connected via wire or wireless network. When wire connection one have to define a network address of computer, for example, 192.168.250.1.

1. To acquire images by DRTECH-panel execute commands: File, Digital devices, Acquire or press button .

2. In the following dialog one can select panel's type and check the panel slot number.

3. Group SSID contains parameters of the wireless connection.

4. To show Serial messages in Log-window check button Show Serial messages.

5. All panel's states are recorded in Status Log window.

6. Exposure time is selected from the corresponding list.

7. The current version of DRTECH SDK has no possibility to acquire by computer command. The acquisition starts by press and hold the button of remote control unit.

8. Press button Start (Calibration) switch panel to accumulate 7 frames for panel's calibration. As in the previous item to acquire the current calibration frame one must press button of remote control unit, or hold it pressed until the 7 frames acquisition. Start button changes to Stop one until end of calibration.

9. After calibration one can save its results in files Gain map (FLAATZ-560.GMP) and Pixel map (FLAATZ-560.MAP), by press the corresponding buttons.

10. FLAATZ-560.GMP and FLAATZ-560.MAP files are saved in the working program folder and are loaded automatically when user selects panel's type (see. item.2).

### VaTech panel (Korea)

To connect Vatech - panel by wire-line network set the network address of computer, for example, 192.168.1.20. The panel address is 192.168.1.80 always.

1. To acquire image from Vatechы execute menu commands: File, Digital devices, Acquire or press button .

2. In the popup dialog select the Path to VADAV.dll.

3. Set the Exposure Time.

4. Make dark calibration by press button Calibration by Dark.

5. Make one or several Bright calibrations by press button Calibration by Bright. Before each calibration one have to switch ON the X-ray radiation for 30%, 60% and 90% of maximum power of X-ray unit (for 3 calibrations). For one calibration the X-ray power must be about 90%.

6. To delete all bright calibrations press button <--Clean. The number of bright calibrations is shown from the left side of button.

7. To acquire an image press button Image Acquire.

8. To correct by dark calibration during acquisition check button Correction Dark.

9. To correct by bright calibration during acquisition check button Correction Bright.

10. To correct the defects of matrix of detectors during acquisition check button Correction Defects of matrix.

11. To rotate image on 90 degree after acquisition select 90° in combo-box Rotation.

12. For mirror transformation of image after acquisition check button Mirror horizontally or Mirror vertically.

13. To crop image after acquisition check button Region Of Interest and set its sizes.

14. The sizes of region of interest can take from the full-size image, which has Region Of Interest, if one press button Get region of interest.

### Select Parameters for Image Acquisition

Program can help to select the exposure time and the number of digital accumulation cycles for the image acquisition by digital devices. To do that:

1. After the execution of commands: File, Digital-device, Select device... check the button Prompt to Exposure Time.

2. Set an arbitrary (the best is small) time of exposure and one digital accumulation cycle in Setup dialog (menu commands: File, Digital device, Setup... or click the button ).

3. Execute the frame acquisition by commands: File, Digital device, Acquire or click the button .

4. On screen appear a new captured frame and the Exposure Time Selection dialog.

5. The window New exposure time shows a time, calculated on base of the acquired image histogram and the old exposure time. You can accept this time without changes or correct it.

6. The window New digital accumulation cycles shows the number of digital accumulation cycles, calculated on base a new exposure time and Acceptable time to acquire one frame. You can accept this value without changes or correct it.

7. In window Acceptable time to acquire one frame set the maximal time, which can be spent for one frame acquire. This time is used to calculate the number of digital accumulation cycles.

8. If you want once more get the prompts for the exposure time and the number of digital cycles - check the button Prompt to exposure time for a next acquire?. If you already have an image with good quality, uncheck this button and close this dialog without parameter's saving.

9. To transfer a new exposure time and a new number of digital accumulation cycles in the digital device acquisition parameters click OK.

10. Execute items 3 - 9 up to get a good quality image.

## Work with TWAIN-compatible Device

TWAIN-compatible devices are devices supporting the TWAIN-protocol of data exchange between the device of input of the raster-type images and computer. Such devices are always sold together with the TWAIN-driver (as a rule it is DLL-file), which is an intermediate between the applied program and device of input of the images. Before start of the program it is necessary to install this driver, following the instructions in a manual to the TWAIN-device.

### Select TWAIN-source

It is necessary to select a type of the device before a work with the TWAIN-compatible device (scanner, digital camera, CCD video-camera etc.) For select a type of the device execute commands: File, TWAIN-scanner, Select Source... .

|  | A dialog to select of one of the Twain-compatible devices, which is registered in the system will appear. Specify a type of the device and click the button Select. The device is selected. |
| --- | --- |

### Image Acquisition by TWAIN-protocol

Execute commands: File, TWAIN-scanner, Acquire or click the button to start of Setup and Acquire dialog of the images by the Twain-compatible device (for example, scanner).

|  | In this dialog you can select halftone modes 8, 10, 12, 16 bits on pixel or color 24, 36 or 48-bit mode - the program automatically converts this image in 16-bit halftone and will load it into a new window. |
| --- | --- |

### Setup Duerr-NDT scanner

To get 16 bits images from Duerr-NDT scanner there must be changed the following files:

C:\WINDOWS\twain_32\vistascan\vsjobtwainmaster.ini and C:\WINDOWS\twain_32\vistaray\vsjobtwainmaster.ini the lines Transfer=MEMORY8 must be replaced to Transfer=MEMORY16.

# Scanning Systems to Acquire Images

## X-ray Scanning System <Orel>

Hardware <Orel> consists from X-ray protected box, which includes: X-ray unit, a positioning system from 5 drives and detector. X-ray unit can move vertically, but a test object and detector have 2 degree of freedom in horizontal planes.

### Prepare Scanning

At first set connection between X-ray control unit and computer:

1. Select type of X-ray unit RAP (Sankt-Petersburg), by execution of menu commands Tools, X-ray Control, Select type.

|  | 2. Call a dialog to control X-ray unit by menu commands Tools,X-ray Control, Setup... or press button . 3. In the following dialog press button PC-control of X-ray. 4. Set Exposure time enough big to our requirements and close the dialog. |
| --- | --- |

Define acquisition parameters for detector tsnkCAM:

1. Execute menu commands Tools, Scan-systems, Orel or press button .

2. Select page Scan.

3. Select type of detector tsnkCAM.

4. Press button Setup scanning system.

|  | 5. In the following dialog select sequentially the frame types: Panoramic view and Region image. 6. For the each type define Exposure time and Black level. 7. Also define names of files to calibrate by black and white images. 8. Detector level must be the same as was defined during assembling of hardware. 9. For the fine tuning of geometrical parameters of <Orel> press button Fine Tuning.... |
| --- | --- |

If system has as detector PaxScan4030 define its acquisition parameters:

1. Execute menu commands Tools, Scan-systems, Orel or press button .

2. Select page Scan.

3. Select type of detector PaxScan4030.

4. Press button Setup scanning system.

5. In the following dialog select sequentially the frame types: Panoramic view and Region image.

6. For the each type define Exposure time and Black level.

7. Angle of the image rotation must correspond to the actual hardware geometry and it should be constant during all operations.

8. Detector level must be the same as was defined during assembling of hardware.

9. For the fine tuning of geometrical parameters of <Orel> press button Fine Tuning....

The dialog of the <Orel> fine tuning should be accessible after user password identification. The parameters of this dialog can change the qualified personal only.

1. For each from 3 formats of panoramic view (1x1, 2x2, 3x2) it should be selected or calculated 6 values: the coordinates of X-ray focus point in template plane and the template's sizes.

2. If the positive axis directions of detector and test object are opposite, then one should check buttons Opposite X-axes or Opposite Y-axes.

3. X-tubeLevel - distance from X-ray focus point to plane of the test-object drives.

4. Level [mm] - a list of three possible distances from the test object plane to the detector plane.

5. Four coordinates are defined a Detector axis ranges.

6. To create the optical correction and joining templates one must check button Create template.

7. And the last parameter - the percent of frame overlap 2x2 for the joining format 2x2.

### Creation and Marking of Panoramic View

1. Fix a test object on the specimen stage.

2. To show the <Orel> control dialog execute menu commands Tools, Scan-Systems, Orel or press button .

|  | 3. Select <Control> page. Press button <Init all drives> to initialize all required 5 drives automatically. |
| --- | --- |
|  | 4. Access to page Setup 1 is possible only after passport input because this page contains a setup parameters of 4 drives to move detector and test object. This parameters should be changed by can change the qualified personal only. There one can save the positions of this drives, which are convenient, for example, to fix the test object near the box's door. To do that all 4 drives move to the required positions and press button Save current position as the loading position. |
|  | 5. Access to page Setup 2 is possible only after passport input because this page contains setup parameters of a drive to move X-ray tube vertically. This parameters should be changed by can change the qualified personal only. There one can set a Movement range of X-ray unit. For example, if you reduce this value, you can test a test objects which are fixed below the usual level safely for X-ray unit. |

6. Check X-ray buttons Switch auto ON and Switch auto OFF to start and finish X-ray automatically during scanning.

7. The format of the panoramic view is depended from the size of test object: for a small objects - 1x1, for big ones - 2x3.

8. Press button Panoramic view.

9. After appear on screen a window with panoramic view press button Create and mark by mouse the interesting regions on this view.

10. To edit positions and sizes of the marked regions one can press button Edit and use mouse. The real size of the region is depends from the defined resolution and format, but the position of region center is defined by mouse.

11. Buttons Save and Load saves and loads a panoramic view together with all parameters and the marked regions. It is helpful for high volume testing of the same objects.

### Frames of Region of Panoramic View

1. Select the number of current marked region.

2. Define resolution to acquire frames.

3. Define format of the region frames.

4. Define level, which have the interesting region.

5. Press button Frame of region to acquire the defined resolution frame automatically.

6. To collect frames of all marked regions of the panoramic view press button All regions. The program moves each region of the test object to X-ray tube and collect acquired frames to a shown on screen sequence. For emergency stop one can use a button .

7. Button Scan movie scans the selected region by frames with defined resolution.

8. To save the collected frame sequence the SET-format is preferable.

### Tomosyntez of Region

To reconstruct the 3D internal structure of the selected region of the test object one can use tomosyntez method. For this method there must be different angle frames of the interesting region, whose can be acquired by the following way.

1. Select the number ofcurrent region.

2. Define resolution to acquire frame.

3. Define format of the region frame 1х1.

4. Define level, where there is a center of object.

5. Define number of projections (frames). It is multiply 4 always.

6. Define the layers of the reconstructed volume.

7. Press button Tomosyntez movie to get the different angle frames with defined resolution. For emergency stop one can use a button .

8. To save the collected frame sequence the SET-format is preferable.

9. To save the coordinates of the marks projections on the start and end layers press button Create tomosyntez marks....

10. This projection data and the marks coordinates are used in tomosyntez method to reconstruct the 3D structure of the test object.

### X-ray Scanning System <Compton-Reverse>

The hardware use X-ray back scattering effect and consists from 2 X-ray tubes (Sankt Peterburg) (new) with rotated collimators and the photoelectronic multiplier detectors moved along the test object. There are 2 controllers to get data from each group of detectors. The controller also controls rotation of corresponding collimator and movement of whole system.

|  | 1. To show the <Compton-Reverse> control dialog execute menu commands Tools, Scan systems, Compton-Reverse or press button . 2. Rotation of collimators for active controllers would be turned on automatically. After dialog close the rotation would be turned off automatically also. 3. Set X-ray parameters of (Sankt-Peterburg) (new) as it is described in this chapter. 4. The current controller can be selected by the corresponding radio-buttons. 5. Time of measurement in point. |
| --- | --- |

6. Buttons Send, Record, Read on can exchange parameters between program and the current controller.

7. Buttons <--, Stop, --> are controls of movement and rotation of the current controller.

8. Button Correct equalizes the brightness of each collimator channels on image in the active window by selected region in the window.

9. Button Start begins image acquisition for the current controller: X-ray unit and motion must be ON. To stop image acquisition press Stop.

10. Button Scan switch on X-ray, start movement along the test object and start data-acquisition for all active controllers.

11. Press button Setup show following popup dialog to setup additional controller's parameters.

12. Best of all there must be defined the controllers to Use for scanning.

13. The number of current controller is defined there also. All parameters are defined only for the current controller.

14. The Setup button shows dialog to select serial port number and its parameters.

15. Each controller collects data from several detectors. The each detector parameters are defined one by one. At first it defies number of detector and then its parameters: PEM-voltage, Detector level, Internal comparator level, Syncronization level.

After that there can be defined the other common for both detector parameters.

16. Select Data collection Mode: Current or Count.

17. If the signal is too weak (less then 128), on can select Bytes per count equal 1. It can help to accelerate data transfer from controllers.

18. Check the buttons Motion-->MDrive and Rotation->MDrive to send in controller the full commands to control movement and rotation, correspondingly. In the other case there are used the short versions of the same commands Left, Stop, Right.

19. To setup serial ports of movement and detector press button Movement and Detector,

20. When Image width autodetection button is checked the program defines image width itself.

21. Otherwise one can set the width himself in the corresponding window.

22. Number of image lines defines length of the ring image buffer.

23. Horizontal axis scale help to set the correct image proportion.

24. Frame offset relative to marker can remove the image distortion.

25. Number of collimators and the position of flags have to correspond to actual hardware configuration.

26. Set Number of averaged lines more than 1 to suppress the image noise, but it increases the acquisition time.

27. Check buttons Correction, Cut+Rotate, Median filter, Scaling to full range to apply the corresponding post-processing automatically.

### X-ray Control Systems

Before usage of X-ray unit it is need to define its type.

1. Execute menu commands Tools, X-ray Control, Select type...

2. Select type of X-ray unit and click OK.

#### Module MGC4X from YXLON

To set the X-ray parameters and to switch ON or OFF the X-ray radiation uses the following commands.

1. Execute menu commands Tools, X-ray control, Setup, or press the button .

2. In this popup dialog press PC-control of X-raybutton to open the serial communication between the X-ray control unit MGC4X and program.

3. If there is some difficulties with open the communication, this button will be unchecked. In this case, please, press the button Serial Port Setup to change the serial port settings.

4. Select Standard or Fine focus of the X-ray tube.

5. Set voltage value [kV].

6. Set value of the current [mA].

7. Press Send parameters button to transfer the corresponding commands to the control unit.

8. The dialog box will be closed automatically.

9. To switch X-ray radiation ON execute the menu commands:Tools, X-ray control, X-ray ON/OFFor press the button .

10. To switch OFF the X-ray radiation repeat item 9 again.

11. To close the serial communication between the program and the X-ray control unit unpress the PC-control of X-ray button.

12. When program is closed it switch OFF the X-ray radiation and close the serial communication with X-ray unit automatically.

#### Module SMART from YXLON

To set the X-ray parameters and to switch ON or OFF the X-ray radiation uses the following commands.

1. Execute menu commands Tools, X-ray control, Setup, or press the button .

2. In this popup dialog press PC-control of X-ray button to open the serial communication between the X-ray control unit SMART and program.

3. If there are some difficulties with open the communication, this button will be unchecked. In this case, please, press the button Serial Port Setup to change the serial port settings.

4. Select Standard or Fine focus of the X-ray tube.

5. Set voltage value [kV].

6. Set value of the current [mA].

7. Select Automatic operational mode of the control unit.

8. To switch X-ray ON or OFF press the buttons X-ray ON! or X-ray OFF!, correspondingly.

9. To switch X-ray radiation ON when the control dialog is closed execute the menu commands Tools, X-ray control, X-ray ON/OFF or press the button .

10. To switch OFF the X-ray radiation repeat item 9 again.

11. To close the serial communication between the program and the X-ray control unit unpress the PC-control of X-ray button.

12. When program is closed it switch OFF the X-ray radiation and close the serial communication with X-ray unit automatically.

#### X-ray Unit (TSNK-lab. MIREA)

To set the X-ray parameters and to switch ON or OFF the X-ray radiation uses the following commands.

1. Execute menu commands Tools, X-ray control, Setup, or press the button .

2. In this popup dialog press PC-control of X-ray button to open the serial communication between the X-ray control unit and program.

3. If there is some difficulties with open the communication, this button will be unchecked. In this case, please, select the type of connection and click the button Connection Setup to set the following serial port settings: 115200 bod, 8 bits data, no parity, 1 stop-bit. In case of Ethernet connection IP-address is usually equal to 192.168.0.30.

4. Set parameters and press button Send parameters.

5. To start X-ray radiation press button Switch ON! or execute menu commands Tools, X-ray control, X-ray ON/OFF or press the button .

6. To switch OFF the X-ray radiation repeat item 5 again.

#### X-ray Unit (Sankt-Petersburg)

To set the X-ray parameters and to switch ON or OFF the X-ray radiation uses the following commands.

1. Execute menu commands Tools, X-ray control, Setup, or press the button .

2. In this popup dialog press PC-control of X-ray button to open the serial communication between the X-ray control unit and program.

3. If there is some difficulties with open the communication, this button will be unchecked. In this case, please, click the button Setup RS-232 to set the following serial port settings: 19200 bod, 8 bits data, parity mark (1), 1 stop-bit.

4. Set parameters and press button Send parameters.

5. To see the working statistics of X-ray unit click Statistics button.

6. To start X-ray radiation press button Switch ON! or execute menu commands Tools, X-ray control, X-ray ON/OFF or press the button .

7. To switch OFF the X-ray radiation repeat item 6 again.

#### X-ray Unit (Sankt-Petersburg) (new)

The program support work with two X-ray units, simultaneously. But its operational parameters would be identically.

To set the X-ray parameters and to switch ON or OFF the X-ray radiation uses the following commands.

1. Execute menu commands Tools, X-ray control, Setup, or press the button .

2. In this popup dialog select the X-ray unit number and press PC-control of X-ray button to open the serial communication between the X-ray control unit and program. The selected unit becomes current one. To operate by another unit one can select another number.

3. If there is some difficulties with open the communication, this button will be unchecked. In this case, please, click the button Setup RS-232 to set the following serial port settings: 19200 bod, 8 bits data, parity mark (1), 1 stop-bit.

4. Set parameters and press button Send parameters.

5. To see the working statistics of X-ray unit click Statistics button.

6. To start X-ray radiation press button Switch ON! or execute menu commands Tools, X-ray control, X-ray ON/OFF or press the button .

7. To switch OFF the X-ray radiation repeat item 6 again.

#### X-ray Unit (Foton, Tomsk)

To set the X-ray parameters and to switch ON or OFF the X-ray radiation uses the following commands.

1. Execute menu commands Tools, X-ray control, Setup, or press the button .

|  | 2. In this popup dialog select a model of X-ray unit from a list. This list has the models which have a corresponding configuration *.RAYfiles in the program folder. 3. Press PC-control of X-ray button to open the serial communication between the X-ray control unit and program. 4. If there is some difficulties with open the communication, this button will be unchecked. In this case, please, click the button Setup RS-232 to set the following serial port settings: 38400 bod, 8 bits data, no parity control, 1 stop-bit. 5. Set parameters and press button Send parameters. 6. To start X-ray radiation press button Switch ON! or execute menu commands Tools, X-ray control, X-ray ON/OFF or press the button . 7. To switch OFF the X-ray radiation repeat item 6 again. 8. To clean information about hardware errors and continue work one can press Stop button. |
| --- | --- |

#### X-ray Focusing Setup

The focusing system of X-ray unit consists from computer controlled drive, which move a constant magnet with ring shape. To setup this system one checks button Focusing and press button Setup... in the setup dialog of the corresponding X-ray unit. All focusing parameters are saved in text file Focus.cor as tree columns: first - X-ray tube voltage, second - current [mkA], third - drive dimensionless coordinate. The number of file lines is not fixed.

1. In the following dialog press button Setup serial port... to define serial communication parameters.

2. Check button Use prefix @, if the corresponding drive is used a new protocol of data transfer.

3. If the previous button is checked then define also one symbol identificator of the serial port.

4. To setup focussing one must select all existing voltages from list one by one.

5. One tunes the values of current and coordinate [impulses] such way that there were sharp images.

6. To estimate sharpness of image one can use menu commands: Tools, Measure, Blur estimation or press button .

To define the optimal focus coordinate for the given values of voltage and current automatically one can use the following procedure:

1. Define a start and final focusing coordinates and a number of steps. The final value must be not bigger than maximal drive coordinate, which can be defined by button Maximal coordinate.

2. Check button Calculate the optimal coordinate.

3. Click button Start. The program automatically create a new window with sequence of frames whose correspond to all focusing coordinates (Those coordinates would be shown in a new window title).

4. Also the program automatically estimates the blur of each frame and calculates the coordinate which correspond to the maximal sharpness and add it into the left side table.

# Image Visualization Methods

The various images (image sequences) are displayed in separate windows inside the basic window of the program. The tool Player allows to control a showing of the various frames from a sequence and their parameters. The name of a current sequence of the frames (or one image) is displayed in title bar of Player. The title bar of window displays number of the frame opening at this moment. It is possible to open new windows for the same sequence to see some frames simultaneously. Moreover it is possible to use 5 alternate methods of visualization in those new windows: Scan, Surface, Tomogram, Insight and 3D-view, which parameters are set by the relevant pages of Player.

## Half-tone Image

This mode of a displaying is the default; i.e. the first loaded image always is displayed as a raster picture.

To scroll the frames of a current sequence necessary to do commands: Tools, Player, or click the button , or type combination of the keys Ctrl+A.

The frame scrolling can be done manually with the help of the slider, or automatically - by pressing of the button Animation. It is possible to set two modes of automatically scrolling of the frames in dialog Options. The relevant buttons of the dialog sets the range and velocity of scrolling.

Frame scrolling is also possible by use the mouse wheel. But if there is Region Of Interest on the image, then to scroll frames one have to press and hold the key Ctrl.

To scroll a visible region of the large image one can use horizontal and vertical scroll bars, or press key Shift and drag image by mouse with the pressed left button.

## Surface View

To display the image as a three-dimensional surface it is necessary to create a window Surface:

1. Show dialog Player by the commands: Tools, Player, or click the button , or click the left mouse button on page Surface.

2. The same effect can be providing with the help of commands: Windows, New Window, Surface view (or button ).

3. The scrolling of surfaces is carried out similarly to scrolling of the raster-type images.

4. The parameters of the surface can be changed in Player dialog on the Surface page and are implemented by pressing of Apply button.

5. It is possible to show a subdomain of all surfaces. For this purpose we set a subdomain in dialog View Options, we click the Use ROI button and then click the Apply button.

6. To select the colors of the text and lines of a surface click Surf. colors button.

7. To cut off a upper part of a surface click the Cut plane button and type its equation.

8. The point of view on a surface varies by two sliders, which located along the lower and right parties of Player.

9. On a surface besides lines of a grid it is possible also to draw level lines if click the Z - contours button and set a range of their change.

## Tomogram View

To show of a sequence of the frames as sequential layers of three-dimensional object it is necessary to open Tomogram window:

1. Open Player dialog by commands: Tools, Player, or click the button , or click the left mouse button on page Tomogram.

|  | 2. It is possible to execute also commands: Windows, New window, Tomogram-view (or button ). 3. The three-dimensional object is dissected by three planes, perpendicular to Cartesian coordinate axes with the help of three sliders X, Y, Z. 4. The point of view on the object varies by two sliders located along the lower and right parties of Player. 5. After removal of a mark from the button Paint all planes over only three secant planes are painted (without lateral areas of object). 6. In this mode it is possible to create "a puff pie" of layers if remove a mark from the button Clean old picture, and moving one of secant planes, and periodically pressing the button Apply. 7. If click the button Show pixel-range and to set them range, on secant planes will be colored only pixels from this range. |
| --- | --- |

## Insight View

To display a sequence of the frames as sequential layers of three-dimensional object it is necessary to open Insight window:

1. Open dialog Player by commands: Tools, Player, or click the button , or click the left mouse button on page Insight.

2. Alternatively, it is possible to execute commands: Windows, New window, Insight-view (button ).

|  | 3. Thus the regions of object contain pixels with values from Min up to Max are considered immersing light with a given Transparency coefficient. Other regions are considered as transparent. Therefore three-dimensional object is shown as the shadow image of structure of regions, immersing light. Than more thickness of region immersing light, especially dark it looks. 4. For the black-and-white image, when the given regions immerse light absolute, it is necessary check the button Black & White. 5. If check the button Outer contour and set its sizes, the view only of this part of volume will be shown. 6. The point of view on object varies by two sliders located along the lower and right parties of Player. 7. The shadow method of visualization of interior structure is most effective in a combination with animation movie, in which the point of view on object varies. |
| --- | --- |

## 3D-View

To display a sequence of the frames as sequential layers of three-dimensional object one can generate isosurface into the test object and visualize it in 3D-view window. To do that there must be installed DirectX version 9.0 or higher.

1. Open dialog Player by commands: Tools, Player, or click the button , or click the left mouse button on page 3D-view.

2. Alternatively, it is possible to execute commands: Windows, New window, 3D-view (button ).

3. Isosurface is calculated for the pixel value Max. (Isosurface calculation for the value Min in this version is not realized).

4. The isosurface detailing is maximal for render step equal 1 - volume-cells is not merged. Value 6 means that approximate 6 volume-cells is merged into the big one, and isosurface is not so detailed. It is allows to greatly reduce the number of triangular surface elements and accelerate the surface calculation.

5. If one check button Use Volume Of Interest and define its sizes, then there would be shown only this selected part.

6. The position of point of view is changed by two sliders which placed along the bottom and right sides of Player.

7. Tree upper sliders are changed the coordinates of the rotation center relative to the object center.

8. Slider Scale changes the distance from the point of view to the test object.

9. Button White background one can use before printout. Otherwise there is used a black background.

10. Check button Light reflection to show the surface light glares from the light source.

11. Check button Light source move to slow circular movement of the light source around the test object.

12. Button Outer contour to show the external object boundary is not realized in this version.

13. This visualization method can be combined with the movie creation, in which the point of view, or/and scale or/and the selected volume can be changed.

## Change Image Scale

To re scaling the image it is possible to take advantage of one of the following methods:

1. On the panel of tools in a window of range select one from values of the list, or type new value and click a key Enter.

2. Execute commands: View, View Options... or click the button . In the appeared dialog to set scale factors separately for width and height.

3. In View Options dialog click the button Fit image to window, then the change of the sizes of a window by the mouse automatically changes also scale of the image, and independent for width and for height.

4. To quick select the scale 100% executes commands: View, Actual View or click the button .

The areas of drawing of a surface, tomograph and insight views can be scaled by the same methods. However, the initial size of these areas is set in the Options dialog.

## Use Two Displays to Image Scaling

The program can works with two displays to scale the images.

For this work it is necessary:

1. Insert into PC the video-card, which support 2 displays.

2. Connect two displays (or display and TV-set) to the video-card.

3. Install the video-card driver.

4. Right click an open area of the desktop and select Properties on the context menu to open the Display Properties dialog box. Click the Settings tab once.

5. Check the button Extend My Windows desktop onto this monitor.

6. Click OK to close the dialog.

When one working with two displays the second one (or on the TV-set) always shows the ROI of the active window on the first display, stretched on the full screen.

Remember, that aspect ratio of the most displays (or TV-sets) is equal 4:3, because that to the keep true image proportions on the second display, one have to select a ROI with the same proportions. It can be done, if select ROI by mouse with the pressed Ctrl key. Otherwise the ROI proportions can be arbitrary and the image pixels on a second display will be not squares.

The ROI sizes can be increased or decreased on 2 pixel by the mouse wheel rotations forward or backward. The ROI center is not changed.

## Scroll through a Big Images

For detailed viewing a big images there exist a special mode.

1. Load a big size image.

2. Execute menu commands Tools, Scroll through image... or press button .

3. In the following dialog define scale of image.

4. Check button On second display, if the second display is connected to the computer. In this case the detailed viewing would be done on the second display, but the first one would show a whole image.

5. Check button Shadow marked parts for shading the looking parts of the image.

6. To change the image regions one can use the navigate buttons in the left bottom corner of the dialog. The image regions are numerated consecutive from the left to the right and from the bottom to the top. The current region number is shown in the right bottom corner of the dialog.

7. If the current image region is tested, it can be marked as ready one by button Done!. During that the next image region becomes the current one automatically. To remove the mark one can use repeated click on button Done!.

8. When the last image region is ready there would be a special warning on the screen.

The scroll mode can be used together with the mode of dynamic comparison of two images. To do that:

1. Load two images.

2. Show Image fitting dialog by menu commands: Tools, Fitting... or press button .

3. There would be the dialog for superposition of the orthogonal transformed superposed image on the current image. The last active image becomes superposed.

4. To select the current image, simply click left mouse button on it and select the region to compare.

5. Check button Start to begin the process of alternative show the superposed and the current images.

6. Define the coefficients of transformation of the superposed image for the coincidence both images.

7. Uncheck button Star to stop comparison.

8. Execute items 2-6 of the previous process to show the Scroll dialog.

9. Check button Start to begin the process of alternative show the superposed and the current images in the selected region of the current image.

10. The navigation buttons of dialog change the selected region of the current image and the comparison is continued on another part of image.

11. When all regions are tested uncheck button Start and close the Image fitting dialog.

## Use Pseudo-colors Palettes

The program has 13 predetermined palettes of colors.

1. The default palette is the palette from 256-halftones grey, which is usually applied to a showing of radiograph. To featuring the infrared images palettes pseudo-color frequently are used.

2. The change of the palette is fulfilled by commands: View, Palette, Palette, clicking of buttons:,,,,,,,,,,,,.

3. It is necessary to mean, that at a loading of the images containing 256 color palette (BMP, GIF, JPEG, PCX, TIFF-formats, 8 bits on pixel) this palette also is loaded and is used for visualization, and colors in the palette reordering by way of increase of their intensity.

4. To save current 256 color palette in a PAL-file (compatible with a format Adobe Photo Shop) or COL-file, execute commands: View, Palette, Load... or click the button .

5. For select and load PAL-file or COL-file with 256 colors palette execute commands: View, Palette, Save... or click the button .

## Use <Look Up and Histogram> tool

The details of the image latent because of small contrast, can be made visual with by the help of operations with LookUp table.

1. For open the LookUp-table menu execute commands: View, LUT or click the button , or type a combination of keys Ctrl+U.

2. The black color here shows a histogram for the current frame, and magenta - diagram of transformation 4096 tones in 256 visual colors of a current palette. In an example is shown, that 256 colors or halftones of the display are used for a showing of intensities from a range from 433 up to 1749. It improves contrast in the given range and erases all details outside of this range.

3. To select the range move it by mouse with the pressed left button, or input the values of low and upper boundaries of the range and press the button Set Range.

4. In the top right corner of this dialog is shown how many percent of the all pixels from ROI are in the selected range of brightness. Evidently, than the less selected range and the more this percent that the better visual perception of the current image (its dynamic range).

5. Button Reset restores the full brightness range.

6. Button Auto selects low and upper boundaries of brightness automatically to show LUT auto-range from Options. dialog

7. The checked button from the left of Auto means that after select a new ROI of image there is a automatic calculation of the optimal brightness range.

8. There exist different types of brightness transformations:

Linear - usual linear transformation of brightness;

Gamma enhances the contrast for the darker parts of images, by the power transformation (with Alpha coefficient);

Inverse gamma enhances the contrast for the lighter parts of images;

Contrast gives the optimal contrast by using the integrated histogram. Sometimes the result can be very noisy and not so natural.

9. Button Negative– invert the current transformation.

## Use <Color Table> tool

1. To change allocation of colors in a current palette use the tool Color Table, which displayed with the help of commands: View, Color Table..., click the button , or type combination of keys CTRL+L.

2. To see current transformation of colors - merely move by the mouse in a central or lower window. Thus in a right window you will see the current number of color, and in the left lower window - relevant number of color in a palette (both values are mapped by red color). The initial color scale is shown in the left-hand vertical window, and used scale - in the lower horizontal window.

3. The button Linear picks the linear table of colors. For this table type you should enter values for the lower and upper color in the left-hand and right lower windows, and then click with the mouse on this button.

4. The button Contrast selects the color table relevant to integrated global histogram. Such table of colors gives the peak contrast for all parts of the image.

5. Button User table selects the table, assigned to the user, of colors. In this mode you have following opportunities:

6. Draw your own table, having pressed the left mouse button and moving it.

7. Select one from the beforehand given tables:

Arcs - table as an arc of a circle,

Steps - table as steps,

Saw - table as tooth of a saw,

Isolines - table as separate peaks.

8. You can execute with the table the following activities:

Copy - to copy current color table into the interior buffer.

Insert - to take the color table from the interior buffer.

Save - to save the table of colors in a file (*.CLR).

Open - to load the table of colors from a file (*.CLR).

9. You can execute operation with the table:

Standard - gives the linear (identical) table of colors,

Negative - creates a negative of the current table of colors,

Logarithm - applies logarithmic function to the current table of colors,

Minimum- takes a minimum from two tables: current and saved in the interior buffer,

Maximum- takes a maximum from two tables,

Multiplication- functional multiplication of two tables.

## Image Sewing from Frames

The program allows making big image from great number of small one. And the small images can beforehand be cut off. The necessity for such operation arises after use of the equipment for scanning extended objects.

1. For piece together of the image from the frames execute commands: View, Frame Gallery... or click the button .

2. Type the number of Frames in line of the new big image.

3. Select Cutted region and its Start. Only these areas from each frame make of the new image.

4. Select one from three of methods to display the frames.

5. During displaying of the frames Line by line all lines of the new image are made from left to right from the frame in ascending order of their numbers.

6. During show of the frame Like snake even lines of the new image are made from the frame in decreasing order.

7. During show of the frame Mixture all columns of the new image are made from above downwards of the frame in ascending order of their numbers.

8. Check button Keep gallery border if you don't want remove the external cutted parts of frames.

9. For return to usual mode of show of the frame (without a "gluing" together) again execute items 1 and 2, but it is necessary to set number of the Frames in a line 0.

The program has two opportunities of the "gluing" together of the frames, but they concern more to numeral processing, than to visualization. The first is Image Gluing, and the second - average with horizontal shift.

# Measurements

## Calibration of the Pixel Values, Length, Time and Angles

1. The menuitem Calibration is used for calibration of measurement values. It is appear with the help of commands: Tools, Measure, Calibration..., or click the button .

|  | 2. This Pixels- page sets a range of pixel values. |
| --- | --- |
|  | 3. The Length- page sets the units and a range of spatial coordinates, and also their initial values. |
|  | 4. The Time page sets of unit and time scale. |
|  | 5. The Angles- page sets the unit and a range of angles. If it is marked the button other, then it is necessary to type a title of measurement unit of angles and number of such units in a full circle (how many such units are necessary on 360° ). 6. Those calibration information can be saved in INF-file by click button Save... and can be loaded by button Load.... |

If in Options dialog button Create info for saved files is checked, then one save an image its calibration information automatically saved in a file with the same name but with extension INF. Later during this image loading its calibration data is loaded also.

## Length Calibration by Standard

Alternatively, if there is an standard of known length on the image, then it is possible to use calibration of spatial coordinates with the help of this standard. It will start on commands: Tools, Measure, Standard or click the button .

|  | It is necessary click the left mouse button at first on one end of the standard, and then on the other, then appear a following dialog, in which are set length of the standard in new units and name of these units. |
| --- | --- |

## Pixel Value Measurement

It is necessary to include a measurement mode for measuring values and coordinates of pixels by commands: Tools, Measure, or click the button , or combination of keys CTRL+M. Then the windows will appear in a status line more to the left of the indicator of time, which showing current coordinates of the cursor and value pixels in a point of the cursor. The cursor can be moved and with the help of the keyboard too. To cancel of measuring mode once again execute the same commands, or depress the mouse button.

## Pixel Values along Lines (Scans)

The Scan page in Player serves for a displaying of allocation of pixel values along horizontal, vertical, arbitrary - oriented of lines, along circles and along a time axis.

1. Open Player by commands: Tools, Player, or click the button or combination of keys Ctrl+A. Click a Scan - page, to activate the Scan-page.

2. Select Scan type and to begin to move the mouse with the pressed left button along the chosen image.

3. The type - Horizontal: switch on a mode of scanning along a horizontal line of the image.

4. The type - Vertical: switch on a mode of scanning along a vertical of the image.

5. The type - Arbitrary: switch on a mode of scanning lengthways of arbitrary oriented line of the image.

6. The type - Depth: switch on a mode of scanning on time in a point of the cursor. This mode shows, how varies the value of pixel in a given point of the image

In this mode you also can spot a reference point on the image. When such point is spotted, the difference between time dependence of a current point and reference point is shown.

7. The type of a Circle - switch on a mode of scanning along an arc of a circle with center in a reference point.

8. The change of scan-lines on the different frames is possible is seen in a hand-operated mode - move upper slider, and in automatically one - click Animate button.

9. The scan diagram can be smoothed with the help of Smooth buttons, which set number of serial smoothing of scan by the formula: Fknew = (Fk-1 + Fk + Fk+1)/3.

10. The scan diagram is possible average and on space in a direction a perpendicular line, for which it is calculated. So if to set Smooth equal to three, for horizontal scan it means, that each point scan is obtained by an average of three quantities: the pixel values are higher a scan-line and values in next pixels and below it.

11. If click the button FFT (Fast Fourier Transformation), instead of the scan diagram the diagram of its distribution of frequencies is drawn.

12. During move of the mouse inside a scan-window the red and dark blue color shows current coordinates of the cursor in this window. And during move of the cursor in a window of the image (together with scan) the red color shows medial value on scan, and dark blue - peak diversion from this medial.

13. Clicking Fix button displays the current image scan by different colors.

14. Clicking Export button save values along scan as the two-column textual table of Microsoft Excel (XLS-file).

15. Clicking Clipboard button copy current scan into the Clipboard as a Windows meta-file.

16. The image of scan can be enlarged by selection of a subdomain, which interesting for you. It can be made by the mouse with the pressed left button or open special menu by clicking of the right mouse button.

## Image Distance Measurement

1. It is necessary to switch on measuring mode for measuring distances by commands: Tools, Measure, Distance & Area, or click the button , or combination of keys CTRL+M.

2. Then it is necessary click the left mouse button on the first point of the image, depress the button and move the cursor to other point. Notice that in a status line will appear the window L:, showing current distance from the first point up to the cursor.

3. To select a new index point - click the right mouse button and execute item 2.

4. For cancel the measuring mode execute item 1.

## Minimal, Mean, Maximal and Dispersion Pixel Value Measurements For the Rectangular Region

To set measurement mode in selected rectangular region one can execute menu commands Tools, Options... and in the appeared dialog check button Statistics for ROI.

To start measurements select by mouse with pressed left button the rectangular region of image. Statistics for this region would be shown in the status bar. This selected region can be shifted when the left mouse button is pressed within the region. The sizes of selected region can be defined in View Options dialog available by menu commands: View, View Options... or by click button .

## Perimeter, Square, Minimal, Mean and Maximal Pixel Value Measurements

The program allows making measurements for polygonal areas. Maximum and minimum values are calculated only for convex areas.

1. It is necessary to switch on measuring mode for measuring distances by commands: click Tools, Measure, Distance & Area, or click the button , or combination of keys CTRL+M.

2. Then to determine points of polygonal area, serial clicking the left mouse button in each point.

3. Notice that in a status line will appear the window L:, displaying current value of perimeter, and window S:, and windows Min, Max, Mean, showing the values of square, minimum, maximal and mean for the area.

4. To begin measuring new area click the right mouse button and execute items 2-3.

5. To cancel the measuring mode execute item 1.

## Reference Point Definition

The reference point is used to create a timegram or scan along a circle on Scan-page of Player.

1. For switch on mode to set on the image a reference point: click Tools, Reference Point, or click the button .

2. Click the left mouse button on the image. If the place of the point does not fit you, then the cursor can be moved to other place and again to click the left mouse button.

3. For cancel the measuring mode execute item.

## Set Marks on Image

The program allows to set on the images up to four (or 14) marks. Three marks are normally used for realization of Affin transformation. Two or four marks are necessary for realization of reconstruction by a Tomosintez method.

1. It is necessary to set of marks: click Tools, Measure, Set Marks... or click the button .

2. In the appeared dialog select number of marks: 4 or 14

3. Check the button Mark 1.

4. Click the left mouse button on the image above the first mark or input its coordinates by hands.

5. Do items 3-4 for a mark 2 (if it is required and for mark 3).

6. Pass to the next frame with the help of slider in the bottom of dialog and to do items 3-5 for all frames of a sequence.

7. Button Find searches the center of mass of most black pixels of the selected region of image automatically. This center becomes the current mark.

8. Button Find All defines more exactly the coordinates of all marks by search the centers of mass of most black pixels of the selected regions around each mark. Size of the all regions is equal to the size of selected one.

9. During arrangement of marks the image is possible to contrast and to change its scale.

10. The coordinates of all marks can be saved in a file, and then quickly loaded from it.

11. Exit dialog to set of marks.

## Measurement angles and conjective distances

To measure angles and distances between the conjugate points on the angle sides:

1. Execute commands: Tools, Measure, Angles... or click the button .

2. Check the button Circle center to define a circle center on the image.

3. Check the button Circumferential points to define a circle center by the points (from 3 to 20) on circumference. Mark those points by the left mouse button and press OK.

4. Check the button Fixed line to define a reference line direction.

5. Check the button Rotated line to define a rotated line direction. The angles are measured from fixed line to this one in the clockwise direction.

6. Check the button Conjugate points to calculate distance between the conjugate points on fixed and rotated lines. Press left mouse button on image - program calculate the conjugate points: the points of intersection of perpendicular to the rotated line, which go through the mouse defined point. If the fixed and rotated lines are perpendicular, then it is impossible define the conjugate points.

7. To show or hide the fixed and rotated lines use the buttons from Show group.

8. To rotate the rotated line type the angle value in the edit-window and press Enter key.

To record the measured angles or distances to a table:

1. Click the button Write to table.

2. Check the button Rotated line or Conjugate points to record the angles or the conjugate point's distances, correspondingly.

3. Press left mouse button to define the angle or the conjugate points, the corresponding value would be recorded to the table. The table can contain up to 20 values.

4. The Clean button deletes all recorded values from the table and hides the empty table.

5. The Paste button pastes all recorded values to the clipboard as text and draw it in the left upper corner of the active window. To change the position of this text one can use the Drawbar tool.

## Statistics for Ring or Round Regions

To make measurements in ring or as particular case in a round region:

1. Execute menu commands Tools, Measure, Ring Region or press button .

2. In the appeared dialog input coordinates of region center or check Center button to select center by left mouse button click on image.

3. Input minimum and maximum radiuses of ring region. For round region the internal radius must be equal 0. Alternatively, one can check Radius button and select radiuses by left button mouse clock on image. To define the radius one can click left mouse button in a point which is near to this radius than to another radius.

4. To define the brightness region one can set minimum and maximum values. Alternatively, one can check Brightness button and automatically select minimum and maximum values from the selected rectangular region.

5. After region's parameters definition one can click Calculate button to find 5 values for the selected region:

a) minimum value,

b) mean value,

c) maximum value,

d) dispersion value (standard deviation),

e) part of area [in percent] occupied by pixels with defined range of brightness.

6. To transfer the calculated values into a text table one can press Copy to Clipboard button, after that a text string consists from 5 numbers, divided by horizontal tab symbol would be in Clipboard.

## Brightness Table

To calibrate the dual energy method there are used stepped standards from metal and organic materials. They have shapes of con-centric rings or parallel stepped surface. To define automatically the mean brightness of each region of standard with the constant thickness one can apply tool Brightness Table.

1. Load an image of standard.

2. Execute menu commands: Tools, Measure, Brightness Table or press button .

3. Select shape of standard.

4. For the rectangular steps shape before item 3 one must select by mouse the rectangular region around the standard.

5. For the con-centric rings shape define the reference point - the standard's center. Internal (R0) and external (R1) radiuses and reference angles from which the number of columns is counted off.

6. Define number of lines and columns in the brightness table.

7. Iterative define the gaps: height and width between lines and columns, correspondingly.

8. Click OK. The resulting brightness matrix is saved in the browsed text file.

## Sharpness Estimation

During optical system installation for sharpness the useful tool is the numerical estimation of the image sharpness.

1. Select rectangular region of image, where one want define the sharpness.

2. Execute menu commands Tools, Measure, Blur estimation... or click button .

3. In the appeared box there would be the value of sharpness. As this value bigger as image is sharper.

## Estimation of Scattered Radiation

As a rule, the scattered radiation is shown as random bright white points on image. Because that to estimate the scattered radiation there is used the percent of white pixels.

1. Select rectangular region of image, where one want estimate scattered radiation.

2. Execute menu commands Tools, Measure, Estimation of scattered radiation.

3. In the appeared box there would be the estimation of scattered radiation. As this value bigger, as there is more scattered radiation.

## Detector Persistence

To estimate the detector persistence there was created hardware with the special version of TSNK-camera to acquire 128 lines from 128 detectors. At first the camera saves 4 lines of black, then switch ON X-ray and after its stabilization save also 4 lines of white. Later the camera switches X-ray OFF and saves 120 lines with exposure time 10 msec.

1. Insert modules of detectors into the corresponding slots.

2. Execute menu commands Tools, Measurement, Detector Persistence.

3. In the appeared dialog select the number of detectors in one module.

4. Check the used modules.

5. Define time after X-ray stop, when the persistence must be measured.

6. Define a text of comment for measurement.

7. Press button Test detectors.

8. On screen there would be a text file with the measurement results for each detector and module.

## Statistics of Detectors

To testing detector's module one can use a measurement tool Statistics of detectors.

1. Insert the testing detector module into slot of the testing hardware.

2. Execute menu commands Tools, Measure, Statistics of detectors....

3. Define number of detectors in a module.

4. Define number of modules.

5. Define the start numbers of detectors in each module.

6. Input additional text information: Detector number and Comment.

7. If during the testing you want change X-ray voltage, then check button X-ray voltage [kV] and fill the corresponding table of values.

8. Click button Dark current to save the mean dark current of each detector. Button Test detectors is activated after that operation only.

9. If button X-ray voltage is not checked, then switch X-ray on by hands from the remote control unit.

10. Click button Test detectors. The results of testing are saved in the working program directory in text file DetectorStat.txt and would be shown on screen.

Each next testing erases the previous results.

# Digital Image Processing

The program has many digital filters and processing methods for improving the quality of the images and noise reduction.

## Region of Interest Selection

Normally operations and filters are applied to all pixels of the frame and to all frames of a sequence.

1. If you want processing only part of the frames, you can take advantage of spin-buttons of Player. It is used for setting of the first and last frame of a range.

2. If you have to process only a part of the frame, you can select this part with the help of the mouse with the pressed left button.

If at that time fix the key Shift, then a selected region has always proportion of the length to the height equal 4:3. It is helpful when the computer is connected to 2 displays and on second one is shown a ROI stretched on full screen.

3. After ROI selection its width and height can be changed on 2 pixels by the mouse wheel. A center of ROI is fixed during this resizing.

4. More precisely working part of the frame can be selected in dialog View Options.

## Image Editing

### Image Resize

|  | 1. Execute commands: Edit, Transformation, Resample... or click the button . 2. Check the button Save aspect ratio to keep the initial proportions of the image. 3. Check the button Interpolate to use the bilinear interpolation at image increasing and average at its decreasing. 4. Set the new sizes and click OK. |
| --- | --- |

### Change the Number and Order of Frames

The operations with change the number of the frames are described in this section work with the complete frames, instead of their parts, even if these parts are chosen.

А) To add the frames at the end of a current sequence of the frames:

1. Load the sequence, which frames should be added.

2. Execute command:Edit, Frames, Add... or click the button .

3. Select name of window with the frames, which will be added at the end of a frame sequence from an active window, click OK.

B) To delete of the frames from a current sequence of the frames:

1. Execute commands: Tools, Player or click the button or key combination Ctrl+A.

2. Select a range of the deleted frames. If it is necessary to delete one frame, make it current.

3. Execute commands: Edit, Frames, Delete... or click the button

C) To add frames of one sequence between the frames of the second sequence, through one:

|  | 1. Load both sequences. 2. Execute commands: Edit, Frames, Mix... or click the button . 3. Select names of windows with the frames, which should be mixed, then press OK. |
| --- | --- |

D) For a clipping each "N"- frame from a current sequence of the frames:

1. Make a sequence of the frames current.

2. Execute commands: Edit, Frames, Select... or click the button .

3. Type the number of the starting frame and which frame to cut out. Press OK.

E) To invert the order of selected range of frames:

1. Select frames to invert order by Player or View parameters dialog.

2. Execute commands: Edit, Frames, Invert or press button .

The first selected frame would be last one and the last - first. The second frame becomes next to last and so on.

### Image Crop

1. Select cut out area of the image by the mouse with the pressed left button.

2. It is possible to execute also commands: View, View options or click the button . Set cut out area in the appeared dialog, then press OK.

3. Execute commands: Edit, Frames, Crop or click the button .

### Image Gluing

A) To glue two frames:

1. Load both images.

2. Execute commands: Edit, Frames, Gluing... or click the button .

3. Select names of windows with the gluing frames.

4. Select type of gluing the frames: in horizontal direction or in vertical direction.

5. Click OK.

B) To glue all frames as gallery:

1. Load image sequence.

2. Execute menu commands: Edit, Frames, Glue Like Gallery... or press button .

3. In the following dialog set the number of frames in line of the compound image.

4. Set cutted region.

5. Set the order of frame's joining:

Like snake - even lines of the new image are made from the frame in decreasing order and odd lines - vice versa.

Line by line - all lines of the new image are made from left to right from the frame in ascending order of their numbers.

Mixture - all columns of the new image are made from above downwards of the frame in ascending order of their numbers.

6. Check button Keep gallery border, if you don't want remove cutted parts on the outside boundaries.

7. Click OK.

C) Besides those glues there is exist Correction and Join of Images for Multi-camera Systems (with optical distortion correction).

### Erase Image Region

1. Execute commands: Edit, Erase, Erase Value... or click the button .

2. Type pixel value for clearing. Press OK.

3. Select erased area of the image by the mouse with the pressed left button.

4. It is possible also to execute commands: View, View Options or click the button . Set cut out area in the appeared menu, then click OK.

5. To clean area - execute commands: Edit, Erase, Erase or click the button or key Delete.

### Pixel Values Scaling

Interior data format of the program - is a integer values from 0 up to 65535 (16 bits). To show pixel values in arbitrary ranges the scale factors are used (the linear transformation is fulfilled). They can be different for the different images . Accordingly, color scale for the different images - is different. For making of the images to one color scale it is possible to execute the following commands:

1. Set boundaries of a blanket scale of pixel values - minimum and maximal values. It can be made having seen minimum and maximum values for each of the images in Calibration.

2. Execute commands: Process, Unary Operations, Scaling Values.

3. Set the values, which found in item 1, in dialog and click OK.

4. Repeat items 2-3 to all images, which should be given in a uniform scale of pixel values.

### Image lines edit

#### Exchange odd and even lines

To exchange even and odd lines of image one must execute menu commands: Edit, Lines, Exchange lines or press button .

#### Shift odd lines

To shift odd lines execute menu commands: Edit, Lines, Shift odd lines or press button . This operation is helpful if the image lines were acquired by probe which is scanning in lines direction and there is exist some mechanical gap odd lines relative even ones.

#### Inverse odd lines

To inverse the odd lines of image execute menu commands: Edit, Lines, Inverse odd lines or press button .

This operation is helpful if the image lines were acquired by probe which is scanning odd and even lines in opposite directions.

#### Sum lines

|  | To sum image lines with its vertical extension and shift one can execute menu commands: Edit, Lines, Sum lines. The operation is useful to increase the resolution of images produced by linear detectors. It increases number of lines and makes sum of those lines with a vertical shift. |
| --- | --- |

#### Exchange half-frames

To exchange half-frames of image execute menu commands: Edit, Lines, Exchange half-frames or press button . This operation can be applied to TV-images, where even and odd lines are saved as separate half-frames.

## Image Transformations

### Image Rotation

To rotate the image on 90° counter-clockwise concerning its center execute commands: Edit, Transform, Rotate, 90° CCW or click the button .

To rotate the image on 90° clockwise concerning its center execute commands: Edit, Transform, Rotate, 90° CW or click the button .

To rotate the image on 180° concerning its center execute commands: Edit, Transform, Rotate, 180° or click the button .

For an image rotation on an arbitrary angle:

|  | 1. Set values, with which the empty areas on a tilted picture, will be filled with the help of commands: Edit, Erase, Erase Value.... 2. Execute commands: Edit, Transform, Rotate, Angle... or click the button . 3. Set quantity of a angle in dialog and direction of rotation. Click OK. |
| --- | --- |

### Mirror Transformation

For Mirror transformation of the image with change by places of the left-hand and right sides, execute commands: Edit, Transform, Rotate, Flip Horizontal or click the button .

For Mirror transformation of the image with change by places of the the upper and lower sides execute commands: Edit, Transform, Rotate, Flip Vertical or click the button

### Orthogonal Transformation

The orthogonal transformation fulfills simultaneously shift, scaling and rotation of the image. Since the transformation is fulfilled in three dimensions, it is possible, for example to change by places time (axis Z) and spatial axes.

|  | 1. Set values, with which the empty areas on the conversed image will be filled by commands: Edit, Erase, Erase Value.... 2. Execute commands: Edit, Transform, Orthogonal... or click the button . 3. Set shifts, scale and angles of rotation in dialog. 4. Set the new three-dimensional sizes of a sequence of the images. 5. Click OK. |
| --- | --- |

### Affin Transformation

The Affin transformation made in the program so that it combines a three marks of one image with the relevant marks of the base image. This operation need for making different the frames to uniform scale.

1. Set values, with which the empty areas on the conversed image will be filled by commands: Edit, Erase, Erase Value....

2. Load both images.

3. Place till 3 marks on everyone.

4. Activate base image.

5. Execute commands: Edit, Transform, Аffin... or click the button.

6. Press Get from window button to read the coordinates of 3 marks of the base image into the corresponding fields. Close the dialog, without pressing the button OK.

7. Make active that image, which marks should be combined with marks of base image.

8. Execute item 5. During this operation the coordinates of 3 marks of the active image(s) are wrote only to the corresponding fields.

9. Click OK.

### Non-Linear Transformation

Unfortunately, in the given version of the program this transformation is not implemented.

### Transformation <Fish Eye>

It makes a cushioning correction of image. It can be done for the brightness values only without geometric transformations.

1.Executes menu commands: Edit, Transform, <Fish-Eye> or press button .

|  | 2. Define coordinates ofcenter of transformation [pixels] 3. Define radius of transformation [pixels]. The radius must be bigger than maximal distance from the transformation center to the image boundary. 4. Check button Transform only image brightness if you want modify the brightness only. |
| --- | --- |

### Trapezoid transformation

It makes compensation of the keystone and semi revolving distortions of image.

|  | 1. Execute menu commands: Edit, Transform, Trapezoid... or press button . 2. Define a keystones dx and dy for top - button and left - right pairs of image sides. 3. Define a semi revolving distortions X and Y for the corresponding pairs of sides 4. Click OK. |
| --- | --- |

### Logarithm of Fourier Spectrum

To get the Fourier logarithm of a spectrum of the chosen area execute commands: Process, Log of Fourier Spectrum or click the button . The new window with the spectrum will appear.

## Background Correction

### Subtract Background Image

For subtraction of the background image from a sequence of the frames:

1. Load the background frame and sequence of the frames.

2. Add the background frame at the end of a sequence of the frames.

3. Make last frame current.

4. Execute commands: Process, Time Operations, Difference or click the button .

5. Now last frame has zero luminosity - becomes black. If it is necessary this frame can be deleted from a sequence.

### Subtract Calculated (Linear or Parabolic) Background

If you haven't the background image, then program can calculate it by a method of least squares and subtract from an original image.

1. If the background will be calculated not on all area, but on its ROI (region of interest), it is necessary to select this ROI by the mouse with the pressed left button, or with the help of commands: View, View Options or click the button . (In the appeared dialog to set "cut out" area and click OK.

2. To delete a linear background (z = A+Bx+Cy) execute commands: Process, Delete Background, Linear or click the button .

3. To delete a parabolic background (z = A+Bx+Cy+Dx2+Exy+Fy2) execute commands: Process, Delete Background, Parabolic or click the button .

4. If on the image the ROI was selected, the question will be: Do you need applying this operation to the image or only to its ROI?

### Black & White Correction

To correct image by black and white calibration images:

1. Load image.

2. Execute menu commands: Process, Delete Background, Black & White... or press button .

|  | 3. In the following dialog press button Black... and select a file with black calibration image. 4. Press button White... and select a file with white calibration image. 5. Click OK. |
| --- | --- |

### Image Equalization

The operation of the image equalization has the same functions, as subtraction of a background.

For example, equalization along horizontal lines means, that mean value and dispersion along any horizontal line of the new image become constant and are equal to given values. A deficiency in this case, is that the transformation will remove horizontal lines on the image. But then it is possible to use equalization along slanting lines, or eight-finger crosses.

|  | For realization of equalization: 1. Execute commands: Process, Equalization. 2. In the dialog select a type Equalization for. For equalization along lines enter value of a slope angle of lines. 3. Set new values Mean and Dispersion. 4. Click OK. |
| --- | --- |

See also Filters of the background equalization.

### Correction of the Linear Detector Images

The images produced by linear detector are characterized by strips due to individual differences of gain and offset of ADC for each detector.

A) The first correction method assumes that on image there are a region corresponding to dark signals of each detector (dark columns), and/or a region without test object (white columns).

1. On image select dark columns by rectangular region. The height of region is not important, because the mean dark values would be calculated for all lines.

2. Execute menu commands Process, Unary Operations, Marked columns equalization....

3. Select type of calibration Black.

4. If it is need, correct the selected column's range.

5. Click OK.

6. Repeat the same procedure for white columns, if they are on image. The order is important there - at first there must be black calibration (subtraction), and then - a white calibration.

7. If each vertical line is corresponded to one detector then select radio-button Column instead Line.

B) The second method doesn’t assume presence of any calibration information on image. It reduces a vertical gradient of image by a least square method. To apply second method on can execute menu commands Process, Unary Operations, Vertical Gradient Minimum.

C) The third method assumes that a image is uniformly for brightness - the mean brightness of each line is approximate a constant. In this case one can use Background correction filters, namely a line equalization.

### Correction of blur

This filter corrects a vertical blur of the images incipient because of duration of transfer of picture lines from the camera <Silar> at a not closed shutter, when the image upper lines is superimposed on underlying one.

It fulfills nonlinear transformation of brightness along columns of the image.

For realization this filter:

1. Execute commands: Process, Filters, Unary Operations, Blur Correction.

2. Set Coefficient of correction - scale factor of superimposition of the upper lines on lower. The more blur, the more should be this coefficient.

3. Click OK.

### PaxScan Corrections

|  | This correction uses a dialog to setup PaxScan 4030 to correct images with sizes 2304x3200. |
| --- | --- |

### Optic Distortion Corrections

To correct optical distortions there is used a reference points method:

1. One makes image of some sample with rectangular grid of points.

2. Then you recognize coordinates of mass centers of the points automatically, and this coordinates are saved in template file.

3. The template is used to correct an image by affin transformation of each triangle of image with vertexes in reference points.

#### Create Optic Correction Template

1. Load image of reference points.

2. Select a rectangular region with those points.

3. Remove background of image.

4. Make an image binarization to extract reference points only.

5. Execute menu commands Process, Unary Operations, Optic Correction, Template....

6. In the following dialog set the sizes of rectangular grid of reference points.

7. Select a type of template: Reference points grid, Polynomial (direct) or Polynomial (inverse).

8. Define a minimal number of pixels in grid point - this parameter is used by algorithm of automatic recognition of reference points.

9. Parameter Cut pixels horizontally set equal to 0. It is used to create templates for multi-camera systems only and it means how many pixels was reduced the grid horizontal size during template creation. This reduction is helpful to correct the convergence of reference points plane and the image registration plane.

10. Press button Template... to define a template file name.

11. Click OK to create template itself.

#### Optic Distortion Corrections

To correct the optical distortions:

1. Load an image.

2. Execute menu commands Process, Unary Operations, Optic Correction, Correction... or press button .

3. In the following dialog press button Template... to select a template file.

4. Select type of template: Reference points grid, Polynomial (direct) or Polynomial (inverse).

5. Frames in line must be 1.

6. Click OK. There would be created a window with a new corrected image.

#### Correction and Join of Images for Multi-camera Systems

To correct and join frames of multi-camera system:

1. Load a frame sequence, produced by multi-camera system.

2. Execute menu commands Process, Unary Operations, Optic Correction, Correction... or press button .

3. In the following dialog press button Template... to select a template file for multi-camera system. It is a text file with a list of optic correction files for the corresponding cameras of multi-camera system.

4. Frames in line must be equal to a number of cameras in one line of multi-camera system.

5. Click OK. There would be created a window with a new composed image.

## Digital Filers

Digital filtration is processing by a method of a "sliding" window, when around of the every changed pixel the window of the fixed size and shapes is selected. Further from all pixels in this window is calculated certain function, which value exchanges value central pixel of a window.

### Set Filter Sizes and Conditions

To select the sizes and boundary conditions of digital filters:

1. Execute commands: Process, Filters, Size & Conditions...

2. In dialog set the width and height of the filter (it are odd always).

3. Check the button Mirror boundary condition, if you want, that for processed pixels close to boundary, window pixels, laying outside of the image, should be completed with mirror mapping. Otherwise for the pixels it will set the default pixel value.

4. Set Threshold value. The operation of a filtration is applied only then, when the calculated value of the pixel differs from current on value of threshold or more.

5. If you have clicked the button Binary mask, it is necessary to select its shape later, having executed commands: Process, Filters, Binary Mask, or click the button .

6. Click OK.

7. The fast alternate method to set the filter size is a window on the Filter-bar. Select in this window the filter size or type new and click Enter.

### Non-rectangular Filter Definition

To define the filters of the not rectangular shape the binary mask serves: a matrix from zeros and units, in which unit correspond to pixels included in the filter window, and zero unused pixels. To set this filters:

1. Execute commands: Process, Filters, Size & Conditions...

2. Check the button Binary mask.

3. Set the Sizes of the filter so that they included width and height of a binary mask. Press OK.

4. Execute commands: Process, Filters, Binary Mask or click the button .

5. Set coefficients of a binary mask. To select the current value use mouse or cursor keys. To change value simply click any symbolical key.

6. To change of all zero values of a mask on unit and on the contrary, click the button Inversion.

7. The binary mask can be saved in a MSK-file and then loaded from it by buttons Save and Load.

8. Click OK.

### Mean

This filter exchanges a central value by mean value on a window. It is elementary flattening filter. It is executed by commands: Process, Filters, Rank Statistics, Mean or click the button .

### Rank Statistics Filters

The filters of rank statistics are based on ordering of all pixel values, which selected in the filter window by their quantity. Then the central value of a window is exchanged with value placed in a particular position in this ranked series.

1. To select the minimum value execute commands: Process, Filters, Rank Statistics, Minimum, or click the button .

2. To select the largest minimum value execute commands: Process, Filters, Rank Statistics, Second Minimum or click the button .

3. To select the median value execute commands: Process, Filters, Rank Statistics, Median or click the button .

4. To select the second-largest peak value execute commands: Process, Filters, Rank Statistics, Submaximumor click the button .

5. To select the maximal value execute commands: Process, Filters, Rank Statistics, Maximum, or click the button .

### Convolution

Convolution is a adding with constant coefficients of all values from a rectangular window of the filter. It is possible to make many known filters with the help of convolution, for example, mean, the unsharp mask, Laplace etc.

For its realization:

1. Execute commands: Process, Filters, Convolution or click the button .

2. Set coefficients of a mask. To select the current value you can use mouse or cursor arrows. To type value press any numeral key or minus. To delete a symbol use the key Backspace.

3. That the brightness of image remained constant Norm should be equal to the sum of all coefficients of the mask. If this sum is equal to zero (for the Laplace filter), then the norm will be equal 1.

4. The typed values can be save and load in the future from MSK-files with the help of buttons: Save and Load.

5. Click OK.

### Image Sharpening

To increase sharpness of the image, it is possible applying the filter of unsharp mask. In this case the central value of a window f is exchanged by value f+C* (f - mean). Where C - coefficient, which is defined in dialog, mean - mean value on the window.

By-effect of this filter is increase of the noise level.

To realize this filter:

1. Execute commands: Process, Filters, Sharpening... or click the button .

2. Set coefficient of sharpness C - then it more, the image will be sharper.

3. Click OK.

### Background Equalization

It is possible equalizing of the background on the image also with the help of filters.

Local Histogram equalizes the images on a local histogram obtained on pixel values in a window of the filter. Thus it makes the image, which has about identical number of pixels with brightness from a range 0 - PixelMax. For good results it is necessary use the filter with big size containing not less PixelMax+1 pixels. Therefore this filter is a very slow one. For realization this filter:

1. Set filter size.

2. Execute commands: Process, Filters, Equalization, Local Histogram.

Background Equalization levels a background, i.e. the central value f of a window is exchanged by value f - mean+(PixelMax+1)/2. Where mean - mean value on a window of the filter. The less size of the filter, the small details will go out on the foreground. For realization this filter:

1. Set filter size.

2. Execute commands: Process, Filters, Equalization, Background.

See also Image Equalization

### Non-linear Filters

#### Roberts Filter

The nonlinear Roberts filter is used for mark of contours on the image and it is determined by the formula:

Fjk = |Fjk - Fj+1k+1| + |Fjk+1 - Fj+1k|

For realization this filter executes commands: Process, Filters, Non-Linear, Roberts.

#### Sobel Filter

The nonlinear Sobel filter is used for mark of contours on the image and it is determined by the formulas:

Fjk = sqrt(X2 + Y2), where

X = (Fj-1k+1 + 2Fjk+1 +Fj+1k+1 ) - (Fj-1k-1 + 2Fjk-1 +Fj+1k-1 ),

Y = (Fj-1k-1 + 2Fj-1k +Fj-1k+1 ) - (Fj+1k-1 + 2Fj+1k +Fj+1k+1 )

For realization this filter executes commands: Process, Filters, Non-Linear, Sobel.

#### Cut Peak Filter

The filter is designed to remove a noise, it selectively cut out sharp deviations (vertical lines and separate points)and not changing the other parts of image. It is applied only in horizontal direction.

For realization this filter:

1. Execute commands: Process, Filters, Non-Linear, Cut peaks or click the button .

2. In the following dialog select one of the filter types: Filter 3x3 or Filter Nx1.

3. Define Threshold and if filter type if Nx1 then select Peak width in pixels. The threshold set amplitude to cut out peaks.

4. Click OK.

This dialog can be used to create a text file with coordinates of rectangular defects, if one press button Create Defect Map... select name of the file and click OK.

#### Bilateral Filter

Bilateral filter is applied to smooth image and preserve sharpness of object's boundaries. It is most convenient to preprocess images of dual energy method.

1. Select ROI on image.

2. Execute commands: Process, Filters, Non-Linear, Bilateral or press button .

3. Select the filter threshold.

4. Click OK.

#### Module of Gradient

The module of gradient filter is used to calculate the absolute value of brightness change in horizontal direction.

1. Select ROI on image.

2. Execute commands: Process, Filters, Non-Linear, Module of Gradient or press button

### User Defined Filter

The program give the user an opportunity to experiment with various filters and create own. It is possible to set the own formula of the filter. This filter will include as variable various statistical estimations calculated on a window of the filter and on additional window. In the formula expression it is possible to use a series of mathematical functions and operations described below. For calculation with the formula the program uses the own interpreter of the formulas, therefore such filtration requires more time, than application of the same filter without the interpreter. To create the own filter:

1. Set filter size.

2. If it is necessary, set a binary mask.

3. Execute commands: Process, Filters, User formula or click the button .

4. If you will use statistical estimations from sub-window, set its sizes.

5. To simplify the calculations it is possible to divide the formula on sub-expression (up to 8). The formulas for this expressions type in the relevant fields.

6. Set a basic formula of the filter in the upper window.

7. The typed formula can be save and load by buttons: Save and Load.

8. Click OK.

Note.

In the formula you can use round brackets, mathematical functions:abs(), exp(), ln(), log(), sqrt(), sign()

|  | Variable |  | Statistics |
| --- | --- | --- | --- |
| f[w] mean[w] sigm[w] med[w] max1[w] max2[w] min1[w] min2[w] leq[w] | Center value Mean Dispersion Median Maximum Submaximum Minimum Second minimum Local equalization | mask nldXY Scal tau | - convolution value - Non-Linear Direction (X,Y=0,...,5) 0=sum, 1=minimum, 2=median, 3=maximum, 4=dispersion, 5=inverse dispersion, - Scalar product - TAU-statistics |
|  | Functions |  | Binary operations |
| pos(x)= eta(x)= LUT(x)= | x for x>0 else 0.1 for x>0 else 0. Look Up Tableа. | +,-,*,/ ^, <, > | - plus, minus, multiply, divide - power, min, max |

Variable are calculated on all values of filtration window in view of a binary mask. If variable has a suffix w it means, that variable is calculated on sub-window.

If in an input field of there is no place or exist iterating expressions, you can determine sub-expressions: k1-k8.

### Frequency Filters

The frequency filters are based on realization of two-dimensional discrete Fourier transformation for deriving the image spectrum. Then the particular frequency band of a spectrum is zeroed, or to it the window function smoothly reducing frequencies in this range is applied. Then an inverse transformation Fourier is fulfilled. In the obtained thus image completely leaves or the given frequency band decreases. This method is especially effective for removal of periodic structures or noise on the image.

|  | For realization this filter: 1. Set the region to process. 2. Execute commands: Process, Filters, FFT Filters or click the button . 3. Select a type of the filter. 4. Set the number of delete frequencies. 5. Set coefficients N and k, which determine window function. 6. Check the button Annulate frequency, for complete resetting of frequencies. However, it can generate high-frequency parasitic oscillations on the image (Gibbs effect). 7. Click OK. |
| --- | --- |

# Image Sequence Processing

## Frame Sequence Too Big for Memory

The program works only with image sequences which can be placed in computer's memory. But for AVI and SET-files with big size there is possibility to visualize those sequences in mode Scan (without timegram) and Surface. For mode Tomo it is possible the visualization Tomosyntez 2 marks (slice) only.

For the just described visualization modes one can use tool Animation to save the visualization results in movie file. Besides that, there is a possibility of digital image processing of the current frame of sequence, but after frame change the results of this processing are not saved.

This limited visualization mode is switched automatically - imperceptibly for user.

## Image Arithmetic Operations

1. To realize the arithmetical operations with the images it is necessary load these images.

2. On each image it is selected the working subarea and a range of the frames, which will participate in operation.

|  | 3. Open Binary operations dialog by the commands: Process, Binary Operation or click the button . 4. Select the type of operation. 5. Set the constants of operation. 6. The images participating in operation are selected. 7. Click the button OK - the result of operation will be placed into the new window. |
| --- | --- |

The quick alternative method to make the binary operation is to execute commands: File, Operation, Binary... or click the button .

|  | In the dialog one have to select by left mouse button the two image files (with the same sizes) and click the button Apply - the result will be loaded into a new window. Evidently, that the type and parameters of the operation must be defined beforehand by the command Binary Operations from the Process menu. |
| --- | --- |

## Subtract Frame from Sequence

This operation is useful to subtraction of a background from a sequence of the frames.

1. Load a sequence of the frames.

2. Select the current frame, which will be subtracted from each frame of a sequence.

3. Execute the commands: Process, Time Operations, Difference or click the button .

As result the current frame becomes zero - black.

4. If the image, which is necessary for subtracting, exists separately from a sequence of the frames, then it should be added in this sequence by the command Add....And then to execute items 1-3.

## Average Frame Sequence

This operation is useful to making the background image or to quick average of the frames.

1. Load a sequence of the frames.

2. Execute the commands: Process, Time Operations, Averaging or click the button .

As result the new window with the new image will appear and everyone its pixel is mean value relevant pixels from each added frames.

## Time Averaging

For each frame of a sequence take "N" (odd number) sequential frames, which close about this frame, (including this frame) and for them calculate average. This mean value replaces the frame. It is analog of the averaging filter only applied not for spatial coordinate, but for the time one.

This operation is useful for noise compensation on the slowly varying frames of a sequence. It is applied for the infrared images.

1. Load frame sequence.

2. Execute the commands: Process, Time Operation, Time Averaging or click the button .

3. Select the number of "N" in the menu and click the button OK.

## Averaging with Horizontal Shift

It averages of the frames of a sequence with constant horizontal shift of each sequential frame. The result is a new frame, everyone pixel which is mean value relevant pixels from each of the summable and shifted frames. The new frame is the longer, the more shift.

This operation is useful to a case of sequential acquisitions of the frames during moving of the camera along extended object.

1. Load frame sequence.

2. Execute the commands: Process, Time Operation, Averaging with Shift , or click the button .

3. Set shift and click the button OK.

## Space Alignment

Space alignment of frames by reference point on each frame executes by menu commands Process, Time Operations, Space Alignment or click button .

# Image Analysis

## Image Characteristics

### Image Comparison

The comparison of the images can be executed by a way:

A) Image Fitting.

1. It is necessary to load both images for this purpose.

2. Execute commands: Tools, Image Fitting or click the button .

3. The menu for combination orthogonally conversed of the image on current will appear. The active image on the moment of occurrence of this dialogue becomes combined.

4. For a select of the current image, press the left mouse button and select area of comparison.

5. The Start button begins/ends the process of change of fitted and current images in the selected area of the current image.

6. Selects coefficients of transformation of the fitted image. They can be used for the subsequent transformation of the fitted image.

B) Realization of an affin transformation of one of the images, for combining three reference points of each image.

### Image Binarization

For transmutation of the halftone image in black-and-white execute commands: Process, Binarization or clock the button .

In the appeared dialog select lower and upper thresholds and click OK.

### Contours Extraction

1. Use the mean filter (commands: Process, Filters, Rank Statistics, Mean or button ) to decrease noise.

2. Use the Sobel filter or Roberts one (commands: Process, Filters, Non-Linear, Sobel or Roberts) for contour extraction.

3. Execute negative transformation of image (commands: Process, Unary Operations, Negative or button ).

4. Execute "skeleton" - transformation (commands: Process, Skeleton or button ). This transformation reduces thickness of black fields up to 1 pixel.

## Recognition of the point objects

Recognition of the point objects is helpful, for example, to recognition of the chips by its pads.

To recognize one need:

1. Load the image of chip and select ROI with pads.

2. Execute commands Tools, Points objects... or press button .

3. In following dialog press button Filtration of objects.

4. Press button Select objects by binarization.

5. Select black and white boundaries by slider.

6. Set minimal and maximal possible diameters of point objects.

7. Press button Recognition of objects.

8. If the most of objects are not recognized, one can change the binarization boundaries and/or unpress button Filtration of objects and unpress and again press button Recognition of objects.

9. To edit positions of objects one can check button Edit objects. Then set mouse cursor near to the editable point and press left mouse button to show menu. To exit from edit-mode uncheck button Edit objects or press right mouse button.

Before points edit it is recommended unpress buttons Filtration of objects and Select objects by binarization, to remove the masking image processing.

10. To show numbers of recognized points check button Show numbers.

11. After recognition of all points its configuration can be saved in data-base TPL-file. To do that one type a comment and press button Save in database....

12. To select folder of database files one press button <...>.

13. To test presence in database our points configuration it is need to define Test of validity [%] and press button Search correspondence in database. Test of validity equal 90% means that the corresponding points of two comprised configurations are on a distance which is less 10% of maximal diameter of points configuration. If in database there are the point’s configurations coincident with current one, then its number and comments are shown in bottom line of the dialog.

14. The comparison of points configurations take into consideration the metrical relations (distances) between points only. The rotations, shifts, mirror transformations and scaling of the points configuration is not an obstacle for the comparison.

15. If the size of the distances are important, then before comparison one must check button Test of size, and before saving of the points configuration in database one must define its real scale (see Calibration of the Pixel Values, Length, Time and Angles).

## BGA testing

BGA testing is similar to the recognition of the point’s configurations. The typical operations for BGA-testing are:

1. Load image and select ROI with BGA.

2. Execute menu commands Tools, Test BGA... or press button

3. In following dialog press button Filtration of points.

4. Press button Select BGA by binarization.

5. Select black and white boundaries by slider.

6. Set minimal and maximal possible diameters of BGA objects.

7. Press button Recognition of BGA.

8. If the most of objects are not recognized, one can change the binarization boundaries and/or unpress button Filtration of points and unpress and again press button Recognition of BGA.

9. To edit positions of objects one can check button Edit objects. Then set mouse cursor near to the editable point and press left mouse button to show menu. To exit from edit-mode uncheck button Edit objects or press right mouse button.

Before point’s edit it is recommended unpress buttons Filtration of points and Select BGA by binarization, to remove the masking image processing.

If database has this BGA, then instead edit one can loads its template from database as described in items 13-14 below.

10. To show numbers of recognized objects check button Numbers.

11. After recognition of all points its configuration can be saved in data-base TPL-file. To do that one type a comment and press button Save BGA-template....

12. To select folder of database files one press button <...>.

13. If database has such template, it can be loaded by button Load BGA-template... .

14. One can also find and load the most suitable template from database by press button Find nearest BGA-template.

15. To execute items 13 and 14 it is very important, that two of the most distant BGA-points was recognized exactly. And it is not necessarily, that all points were recognized. The comparison of BGA-points takes into consideration the metrical relations (distances) between points only. The rotations, shifts, mirror transformations and scaling of the BGA-points are not an obstacle for the comparison.

16. To estimate a BGA-quality it is need to define Test of validity [%] and press button Find nearest BGA-template. Test of validity equal 96% means that the BGA-point is OK, if black pixels consists 96% or more all pixels within circle with maximal diameter (50 pixels in this case). Those points are drawn by blue color but defect points - by red color. Besides that the total number of bad points is shown in information line of the dialog.

# Material's Recognition by dual energy method

The dual energy method is based on the different absorption of X-ray energy by different materials. It is an estimation of two material's thicknesses by two X-ray images for two energies of radiation. When those thicknesses are defined in the each point of images one can use pseudo-colors to show regions with different materials. The most interesting is a dividing of regions with metal and organics on the image.

## Dual Energy Method Calibration

Automatic calibration of dual energy method is realized for 2 types of X-ray detectors: FPD PaxScan4030 from Thales and for tsnkCam.

1. Select type of X-ray unit (for example Smart of Yxlon) and run a parameter's dialog by menu commands Tools, X-ray Control, Setup... or click button .

2. Click button PC-control of X-ray and close the dialog.

3. Execute menu commands Tools, Matherial's Recognition, Acquire Images or click button .

4. In popup dialog click button Table to select corresponding file.

5. Additional parameters to setup Dual

|  | Energy method can be defined after check the button Setup and password input. Those parameters can be changed by qualified personal only. 6. Click button Calibration. 7. There would be warning to remove a test object. Then program automatically produce a calibration: acquire black frame and two frames for different energies. Calibration files are in the same folder, where is a dual energy table file. |
| --- | --- |

## Acquire Two Images of Dual Energies

The automatic acquisition of two dual energy frames is realized for 2 types of X-ray detectors: FPD PaxScan4030 from Thales and for tsnkCam.

1. Select type of X-ray unit (for example Smart of Yxlon) and run a parameter's dialog by menu commands Tools, X-ray Control, Setup... or click button .

2. Click button PC-control of X-ray and close the dialog.

3. Execute menu commands Tools, Matherial's Recognition, Acquire Images or click button .

4. In popup dialog click button Table to select corresponding file.

5. Click button Start.

|  | 6. The program acquires frames for two different energies automatically and show a dialog Material's recognition. |
| --- | --- |

## Metal and organics thicknesses estimation

To work in dual energy mode it is necessary do the following operations.

1. Load two images, acquired for different energies.

2. It is desirable to make digital processing of those images by bilateral filter.

3. Execute a menu commands Tools, Dual Energy, Estimate thicknesses... or press button .

4. In nascent dialog select the names of the corresponding images.

5. By press button Table... select a table filename for the same energies.

6. Check the button Windows of organics and metal thicknesses to show two additional windows.

7. Check the button Clusterization and define its parameters to use a clusterization algorithm for post image processing.

8. Press the button Estimate to calculate thicknesses of metal and organics и to show a color image. The regions with prevalence of metal are painted by blue color, and ones with prevalence of organics - by green color. The regions where the material is not identified due to its big thickness are painted by orange-black colors.

9. Priority, brightness and contrast of metal and organics regions can be changed separately by the corresponding sliders after calculation.

The clusterization parameters can't be changed after calculation.

To show an image, produced by the dual energy method it is used a special dual palette. There is no sense to use some other palettes for this image.

## Create a dual energy table

|  | The interpolated table of the dual energy method is created for the selected pair of energies with use of the special stepped tests from metal and organics by the following way. 1. Acquire the images of tests for each energy. 2. By use LUT tool calculate a brightness tables, corresponding to the different combinations of metal and organics thicknesses for each energy. 3. The coefficients of the polynomial approximation of the brightness logarithm for each energy are calculated by least square method. 4. Calculate the table itself by menu commands Tools, Dual energy, Dual energy table... or press the button . At this dialog one use the previous calculated coefficients. Also we define a maximal metal and organics thicknesses and press OK. |
| --- | --- |

# Tomosyntez

The reconstruction by a tomosyntez method counts a current sequence of the frames by a sequence of projections with the different focus and synthesizes from them the interior structure of three-dimensional object.

## Use 2 Image Marks for Tomosyntez in Non-stationary Conditions

It is supposed, that the geometry of deriving of the projection images precisely is not known. However at deriving a projection date the planes of travel of object and source should be parallel to plane of recording. It are supposed rotating of test object only around of an axis of a perpendicular plane of recording.

Put on the test object 2 marks (made from X-ray contrast material) which should be good visible on all projections.

The first mark should be on the lower surface of object, and the second - on the upper surface. And both marks have to be far one from another in horizontal direction.

Because of the precision of tomosyntez method is depend to the precision of measurement of those marks centers, that the most convenient shape is a cross from the thin tungsten wire.

Then for realization of tomosyntez:

1. Load a sequence of the projection images.

2. Open dialog to set marks by commands: Tools, Set Marks or click button .

3. Set 2 marks on each frame (projection). The first mark should display a image of a point on the lower surface of object, and the second - similar image of a point on the upper surface.

4. Open Player by commands: Tools, Player or click button , or press combination of keys Ctrl+A. Click Tomogram-page.

5. Select the Visualization method Tomosyntez: 2 marks (3D) or Tomosyntez 2 marks (slice) to open menu with parameters for tomosyntez.

6. Set the geometrical sizes of test object.

7. Select operation with projections to process projections. The most useful operations are: Average, Submaximum and Second minimum.

8. Click OK.

9. After these operations the Tomogram-page of Player displayed not the initial projection data, but reconstructed interior structure of the object.

For the visualization method Tomosyntez 2 marks (slice) there is reconstructed only one slice of the test object. The depth of this layer is defined by Z-slider and the inclination angles on X and Y - by vertical and horizontal sliders, correspondingly. X and Y sliders is not working in this case.

10. To change the geometry parameters or method to sum projection iterate items 5-7.

## Use 4 Image Marks for Tomosyntez in Non-stationary Conditions

It is supposed, that the geometry of deriving of the projection images precisely is not known.

Put on the test object 4 marks (made from X-ray contrast material) which should be good visible on all projections. The 3 mark should be on the lower surface of object, and 1 - on the upper surface. (If you set 1 mark on the low plane and 3 on the upper one - the reconstructed image would be inverted.)

Because of the precision of tomosyntez method is depend to the precision of measurement of those marks centers, that the most convenient shape is a cross from the thin tungsten wire.

Then for realization of tomosyntez:

1. Load a sequence of the projection images.

2. Open dialog to set marks by commands: Tools, Set Marks or click button .

3. Set 4 marks on each frame (projection). The first mark should display a image of a point on the upper surface of object, and other - similar images of points on the lower surface. It is possible to use 3 marks on upper plane - see remarks at the beginning of this section.

4. Open Player by commands: Tools, Player or click button , or press combination of keys Ctrl+A. Click Tomogram-page.

5. Select the Visualization method Tomosyntez 4 marks to open dialog with parameters for tomosyntez.

6. Set the geometrical sizes of test object.

7. Select an operation with projections to process projections. The most useful operations are: Average, Submaximum and Second minimum.

8. Click OK.

9. After these operations the Tomogram-page of Player displayed not the initial projection data, but reconstructed interior structure of the test object.

10. To change the geometry parameters or method to sum projection iterate items 5-7.

## Correction of the Reconstructed Layers for the Cone Beam

For the cone beam of X-ray radiation the reconstructed layers have different scales: as more distance from X-ray focus point as bigger. To reduce all layers to the same scale one can do the following:

|  | 1. Execute menu commands Process, Time Operations, Tomosyntez correction.... 2. Define coordinates of the projection center and its height above the middle layer (frame) in pixel coordinates. 3. Click OK. |
| --- | --- |

## CT - reconstruction

The classical CT-reconstruction by back-projection method is realized in the VIAM-tomography system. That means this system give a CT-reconstructed slices immediately after scanning of the test object.

## 2D CT for the Classical Projection Data

To execute 2D computer tomography for the acquired data:

1. Execute menu commands: File, Operations, 2D CT..., or click button .

2. In the appeared dialog click button Raw... to browse filename of the projection data.

3. Click button CT... to browse filename of the resulting file.

4. Select CT scan mode.

5. Select reconstruction filter обработки проекций.

6. Select CT image size and projection views.

7. Select rotation direction during acquire projection.

8. Define reconstruction parameters: DSD - distance Source-Detector, DST - distance Source-Test Object, center shift - offset of rotation center relative to center of the projection data and so on.

9. Check the additional functions of processing.

10. Click button Start.

## Pseudo 3D-CT

To make pseudo 3-dimentional computer tomography (3DCT) reconstruction, from the previously acquired data:

1. Load one of the projections to select the region of interest on it.

2. Execute the following menu commands: File, Operations, 3D CT..., or press button .

3. In this popup dialog press Take from... button to browse AVI or BIN file with the projection data. If there was selected BIN-file then the program automatically extracts from its name the base name of projections. Program automatically extracts the projection's base name. (Please, don't use in the filenames the point symbols!) The corresponding base projection's name will be shown at the left side of this button.

4. Define Start number and Number of projections.

5. Press Get sizes from View button to automatically get from the active window the projection's Width and Height, and also Topmost,Total number and Center of rotation from the ROI on this active window. Increment is equal 1 always after this operation.

6. Correct the Center of rotation, Increment and Total Number of slices values, which was defined automatically.

7. Define angle of the full object's turn.

8. Define Slice size from the combo-box.

9. Check Save in one file button to save all resulting slices in one file. Otherwise they will be saved in a separate files with names produced from the base slice name and the slice pixel position from the top of projection data.

10. Check Show after CT button to automatically loading the calculated slice data files after its reconstruction.

11. Press Save as... button to select the base name of the slice(s). This name will be shown from the left side of the button.

12. Check Use filtering button to apply a convolution with the filtering function to the projection data.

13. Check Logarithm correction method to apply a logarithm correction for the projection data.

14. To apply a logarithmic correction you must also define a number of reference points (number of points on the left and the right side of each projection line) to calculate a reference value.

15. Select CT-scan mode in the combo-box. The current version of program supports only Parallel mode.

16. Click OK button.

# Prepare Documentation

## Drawing on Images

For drawing on the image execute commands: View, Toolbars, Drawing or click the button .

The panel of drawing will open which includes the buttons for drawing the elementary geometrical figures, text and for insert raster-type images from buffers of an exchange.

1. To draw a pointer:

a) Click the buttons relevant to the pointer and its color on the panel of drawing.

b) The mouse cursor will change the shape now it is possible to click the left mouse button in the necessary place. The small pointer will be drawn atop of the image.

c) To change its sizes and orientation click the left upper button of the panel.

d) Click the left mouse button on the pointer and around of the pointer will appear the rectangle with the changed sizes. Change the sizes.

e) Click the left mouse button outside of the rectangle and it will disappear. The drawing is finished.

2. At drawing of a text there is the button (in the bottom of the screen) to select size of a font and type of the text, and also field for input of the text. The text orientation can be changed, resizing a rectangle of text editing in appropriate way.

3. For drawing raster-type picture at first it should be inserted into the clipboard with the help of any pictorial editor or load picture in the program and execute the commands: Edit, Copy (or click the button ). The picture sizes and its orientation can be changed, by stretching a rectangle of picture editing in appropriate way.

For removal of the separate drawn object it is necessary:

a) Click the left upper button of the panel of drawing.

b) Click the left mouse button on the object. Around of the object will appear the rectangle with the changed sizes.

c) Click a key Delete on the keyboard.

To hide figure it is possible to remove the mark from button Show overlay in dialog View options.

4. To make in the upper left-hand corner of the frame a three lines with the textual information execute commands: Tools, Image Specification or click the button .

|  | The dialog to input those three test lines of the comment will appear. In this dialog it is possible also to select the size of the font. To remove or change a position of this comment execute commands: View, Toolbars, Drawbar or click the button . It is possible to change names of lines of the comment in dialog Options. |
| --- | --- |

5. During saving of the image in format of 8 bits on pixel the drawing together with the image is conversed to the raster-type image and is saved. During saving of the image in the other format, figure is saved in a separate WMF-file in the same folder and with the same name, but with file name extension WMF.

At reading any image in its folder the file of drawing with a name of the image and expansion WMF is searched. If such file exists, then it will be read out and place on the image.

## Save Image

To save image from the active window execute commands: File, Save as... or click the button or the key combination Ctrl+S.

Program supports formats: AVI, BMP, B16, BIN, FLC, FLI, FITS, GIF, JPEG, PCX, PNG, RAW, SET, TIFF, TXT, XRD, XYZ.

If there was a drawing on the image, then for the image saving with 8 bits per pixel, the drawing is saved as half tone image. If one selects another file format, the drawing is saved separately as WMF-file in the same folder and with the same name, as image.

## Copy Images to Clipboard

One can use the clipboard to transfer 8 bits per pixel images to a text processor or a another graphical software.

1. Execute commands: Edit, Copy or click the button or the key combination Ctrl+C to copy the image from the active window to the clipboard.

2. To paste a scan-line into your report, click the button Clipboard in Player. After that use command Paste from Clipboard in your text processor.

## Print Image

To print image from the active window execute commands: File, Print or click the button or the key combination Ctrl+P. Before the image printing there would be a dialog to select printer and to setup its parameter.

## Create Demo Movies

There are two modes to create animation movie (AVI, FLC or FLI- formats) – manual and automatic. In the manual mode you create a new picture and add it to the movie as one frame. Then you create a next picture and add it and so on.

In the automatic mode at first you create a movie scenario from the nodal pictures and the numbers of frames between those nodes. Then program automatically calculate all intermediate frames and compress to the animation file.

Note, that all images, which is added to a movie have to be created in the same window.

А) To create animation in manual mode:

1. Activate a window to create a movie frames.

2. Execute commands Tools, Animation or click the button .

3. Check the button Manually by one frame in the appeared dialog.

4. Define the movie name in File Open dialog.

5. Set delay time between frames.

6. Create a first frame in the active window.

7. Set the number of the movie frames. Typically it must be 1. But if you want that some frame was repeated several times set a bigger number.

8. Press Add frames button. If you add a first frame to the AVI-file there must be Compression dialog to define the method and the quality of video-compression.

9. Repeat items 6-8 for each movie frame.

10. If there was your error press the Reset scenario button to delete a created scenario and start it again.

11. Click the button Close movie or close the dialog.

Б) To create animation in automatic mode:

1. Activate a window to create a movie frames.

2. Execute commands Tools, Animation or click the button .

3. Check the button Automatically in the appeared dialog.

4. Set delay time between frames.

5. Create a first frame in the active window.

6. Click the button Add frames to add the first node to the movie scenario.

7. Create second nodal frame. Difference between nodal frames must be in the values of the following 9 parameters of visualization:

The frame number (for the all view types),

The horizontal and vertical angles of the view orientation (for Surface, Tomogram and Insight views),

X, Y, Z slice positions (for Tomogram view),

The low and upper boundary of the pixel values (for Insight view),

The transparency coefficient (for Insight view).

If all this parameters are identical for the two nodes then all intermediate frames are also identical, because program does the linear interpolation of this parameters for the intermediate frames.

8. Set the number of intermediate frames between first and second nodal frames of scenario.

9. Press Add frames button to add the current node of scenario.

10. Repeat items 7-9 for each nodal frame of the scenario.

11. If there was your error press Reset scenario button to delete created animation scenario and start it again.

12. Press Create file button.

13. Define the movie name in File Open dialog. If you define AVI-file there must be Compression dialog to define the method and the quality of video-compression.

14. Wait some time while program calculate all intermediate frames and compress it to the movie. Press Stop button if you want to break this process.

# Customize Program to User

## Customize User Interface

The customization of use interface includes the change of shape and contents of toolbars and program menus to adapt it to the user. All user interface parameters are saved in Windows system register for each user separately.

### Customize Commands

To setup tool bar's buttons and menu items:

1. Execute commands: View, Toolbars, Customize... or click the right mouse button on the menu bar or any other toolbar and in the popup menu select the bottom menu item.

2. Select the page Commands.

3. After that there is a mode to delete, add or change the tool bar's buttons and the menu items.

4. In Categories box select one of the main submenus of the program. During that time in the Commands list the all commands available by default for this submenu will be shown.

5. Those commands one can drug by mouse with the pressed left button to any place in any menu or any toolbar.

6. Also it is possible to exchange the places of buttons and menu items.

7. To delete a command from the menu or toolbar click the left mouse button on it and drug it from menu or toolbar.

8. Click the button Close.

### Customize Menu

|  | To setup menu: 1. Execute commands: View, Toolbars, Customize... or click the right mouse button on the menu bar or any other toolbar and in the popup menu select the bottom menu item. 2. Select the page Menu. 3. Select Main menu or Default menu of application or context menu. 4. Click the button Reset to restore the default form of those menu. 5. Select one from the Menu animations. 6. To show the menu shadows check the button Menu shadows. 7. Click the button Close. |
| --- | --- |

### Customize Toolbars

To setup the toolbars:

1. Execute commands: View, Toolbars, Customize... or click the right mouse button on the menu bar or any other toolbar and in the popup menu select the bottom menu item.

2. Select the page Toolbars.

3. Click the button New... to create a new toolbar, and input its name.

4. Or select a toolbar to edit.

5. Click the button Rename... to change a name of the selected toolbar.

6. Click the button Reset to restore on the selected toolbar the all default buttons.

7. Click the button Reset All to restore the default buttons on all toolbars and all default menu items in all menu bars.

8. Click the button Delete to delete the selected toolbar.

9. Check the button Show text labels to show text labels under its icons on the selected toolbar. After that the size of buttons is increased dramatically.

### Add User Tools

|  | To add User Tools: 1. Execute commands: View, Toolbars, Customize... or click the right mouse button on the menu bar or any other toolbar and in the popup menu select the bottom menu item. 2. Select the page Tools. 3. Click the button Insert. 4. Input a name of new menu item: 5. Click the button "..." to browse a tool on disk or input its name directly. 6. Input a possible tool's arguments. 7. Set the name of initial directory. To delete a tool repeat items 1 and 2. Select a tool's name in a list and click the button Delete. |
| --- | --- |

### Customize Keyboard Keys

|  | To setup "hot keys" for the menu commands: 1. Execute commands: View, Toolbars, Customize... or click the right mouse button on the menu bar or any other toolbar and in the popup menu select the bottom menu item. 2. Select the page Keyboard. 3. Select for what menu you want to set Accelerator for. 4. Select the command's category. 5. Select the command to set key. 6. Press a new keys combination to assign for this command. 7. Click the button Close. |
| --- | --- |

### Customize Double Mouse Click

|  | To setup the double mouse click: 1. Execute commands: View, Toolbars, Customize... or click the right mouse button on the menu bar or any other toolbar and in the popup menu select the bottom menu item. 2. Select the page Mouse. 3. Check the button Associate double click event with. 4. Select a name of command, that would be executed after double mouse click on Image window. 5. Click the button Close. |
| --- | --- |

### Customize Styles of Toolbars and Menus

|  | To change the performance of toolbars and menus: 1. Execute commands: View, Toolbars, Customize... or click the right mouse button on the menu bar or any other toolbar and in the popup menu select the bottom menu item. 2. Select the page Options. 3. Check those buttons, whose are need for you. 4. Click the button Reset my usage data to restore the all default menus and toolbars. 5. Click Close button. |
| --- | --- |

## Create and Use Macro-Commands

To execute a routine operations automatically one can use a macro-commands. Macro-command is a sequence of usual menu commands (or button's pushes), which was created and saved by user, and can be started to execute by press one button (or the menu item's execution).

Program has memory for 10 macro-commands, each can contain up to 100 usual commands. All macro-commands are saved in file DIADA.MAC. The current program's version don't support saving the command's parameters (all values which user can set before the command execution) together with macro-command.

To create (edit) macro-command:

1. Execute the following command sequence: Tools, Macro, Edit or click the button .

|  | 2. After that this dialog is appeared and program begin recording of button's press on toolbars and menu items. All image operations a executed as usual. 3. Select Number of macro-command to edit. 4. To clean a list of commands of the current macro click the button Clean. 5. To insert a new command into a command list execute the corresponding command by click a button on toolbar or by select a menu item. A new command would be added into a command list after the selected command from the list. 6. To move the selection through a list use mouse or cursor keys. A short description of the selected command is shown in the button window of the dialog. 7. To stop the command's recording temporary - fix the button Record pause. To resume the command's recording - unfix this button. 8. To delete the selected command from the list click the DELETE key on keyboard. 9. Close dialog by execute item 1 or by click the upper right button of the dialog. |
| --- | --- |

To execute macro-commands 1-10:

1. Execute commands Tools, Macro, Command 1-10 or click the on of the buttons: ,,,,,,,,,.

2. During the macro execution there is showing the progress bar and the dialogs to input the necessary command's parameters. If the program was run in the Foolproof mode, then the number of those dialogs reduced to essential minimum.

3. To stop execution of the longtime macro command click the button or key Esc.

## Fool-Proof Protection

The foolproof protection includes:

1. Customization of the menu and toolbars to delete unused commands and to simplify the operations with usable commands.

2. The password protection of the possibility to change menus and toolbars.

3. Creation of the macro-commands, to execute a complicated operations automatically.

4. Save in Windows system register for the each user the all values of the dialog's parameters (the default parameters). This option simplifies the dialog parameters changing.

5. The execution of commands with the default parameters, without dialogs to input parameters, everywhere, where it is possible. For example, it is impossible to define a filename in File Open dialog beforehand.

To switch ON the protection mode (items 2, 5):

1. Execute commands Tools, Options. or click the button .

2. In the dialog check button Password protected options.

3. Press OK.

To switch OFF the protection mode:

1. Execute commands Tools, Options. or click the button .

2. In the dialog uncheck the button Password protected options.

3. Press OK.

4. In the Input password dialog type the password and click OK of this dialog.

5. If there was typed the correct password the protection would be removed. If the password was incorrect, it is appear the warning "Incorrect password". One can retype the password, starting from the item 3, or close the Options dialog without the protection removal.

# File Formats

#### AIM

AIM-file format contains a binary uncompressed 16-bit-image sequence.

#### AVI

AVI-file format is a standard Microsoft Windows format for the sequence of color or half-tone 8-bits images.

#### B16

B16-file format contains a binary uncompressed 16-bits image. It was created by a German PCO firm.

#### BIN

BIN-file format was created in Beijing University of Aeronautics and Astronautics (China). It is format for the sequence of frames. Each pixel of frame is a floating point value. The pixels are stored in usual order: first frame top line, ..., first frame last line, second frame, ..., last frame. After pixel's data there is a header with the 512 bytes length. From this bytes are used now only last 20 bytes. Bytes 493-496 and 497-500 - floating point values of minimal and maximal pixels in all data. Bytes 501-504, 505-508 and 509-512 - are integer values of width, height and number of frames, correspondingly. The first bytes of file header are saved geometry information about VIAM-tomograph.

#### BMP

Microsoft Windows Bitmap (BMP) file format is one of several supported by Microsoft Windows.

Program converts BMP-file to 8 bits half-tone image. But Preview button of the File-Open dialog give you possibility to see all Windows bitmap files in color.

#### DTV

DTV-format was created by IRTIS developing group in Moscow. This format saves infrared image sequences. The data has 8 bits per pixel and file header 110 bytes.

#### FITS

Flexible Image Transport System is a general-purpose data storage format used primarily for the interchange of data between hardware platforms and software applications. It is standard format for many astronomical organizations, including the astrophysics branch of NASA. The image data are stored uncompressed with unlimited gray scale.

#### FLC

FLC- file format was created by AutoDesk for 8-bits image sequences. It save the RLE - compressed data without loose and give good results for the artificial CAD-drawing bitmaps.

#### FLI

FLI-file format was created by AutoDesk for 8-bits 320x200-image sequences. It save the RLE - compressed data without loose and give good results for the artificial CAD-drawing bitmaps.

#### GIF

Graphics Interchange Format is a creation of CompuServe Inc., and is used to store multiple bitmap images in a single file for exchange between platform and systems. Program supports only 8 bits per pixel images in interlaced or non-interlaced formats and also so-named animated format which can keep an image sequence.

#### JPEG

JPEG-File Interchange Format is a development of C-Cube Microsystems for the purpose of storing JPEG - coded data. This version of program supports 8 and 12 bits per pixel gray and color images of this format.

#### PCX

PCX-file format originated with Zsoft’s MS-DOS-based PC Paintbrush. Program supports only 8 and 24 bit per pixel, compressed PCX-images.

#### PNG

Portable Network Graphics (PNG), for a lossless, portable, compressed individual computer graphics image transmitted across the Internet. It is an extension of GIF - format. Indexed-color, grey-scale, and true-color images are supported, with optional transparency. Sample depths range from 1 to 16 bits.

Program support read 8, 16 and 24 bits per pixel images and write 16 and 8-bits images.

#### RAW

RAW-file format is binary format for which one point of data can have C-type byte, int, word, long, double word, float or double. You must define the type of values, width, height of frames, depth or number of frames and the byte’s offset of values array in the file in the dialog box.

#### SET

SET-file format is ASCII-format for an image sequence. It was created by Chakhlov Sergej (Introscopy Institute, Tomsk, Russia). The first three file lines are number of frames, width and height of frames. Next lines contain information about each frame: offset of frame’s data in file and filename. Last two lines in SET-file are floating point values corresponding to the first frame and to the second one. For example, they can be numbers of frames, time’s values of frames or depth’s values. Data in each frame-file must have a same size. Supported formats for frames are BMP, B16, DTV, GIF, JPG, PCX, RAW, TIFF.

#### TIFF

TIFF-format is perhaps the most versatile and diverse bitmap format in existence. Its extensible nature and support for numerous data compression schemes allows developers to customize the TIFF format to fit any perculiar data storage needs. Program supports only 8, 10, 12, 14 and 16-bits per pixel images uncompressed or compressed by LZW -algorithm. For 8 bits it supports also color images.

#### TXT

TXT - is ASCII-format to save one image as a text table, in which the tab symbol is a separator of a pixel values in a line. The carriage return symbol is a line separator. The pixel values itself can be integer or floating point values. This format is easy to import to the Excel.

#### XRD

XRD - format was created by TSNK-lab MIREA (Moscow, Russia) to save a binary data (ZIP-compressed) produced by a scanner with linear detectors. Program supports read and write format.

#### XYZ

XYZ - is a binary two-bytes format with header 4 bytes contains width and height of the image. It was created by German firm Duerr-NDT. Before the image data there are two 2-bytes values of image width and height. Program supports read and write format.